| Item 1 | Gynaecology & Hysteroscopy Surgical laparoscopic Tower, Full HD | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |
| --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |
| FDA clearance OR CE Mark |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | Durable, heavy duty and reliable from a well-known manufacturer |  |  |
| A | Full HD Camera head: |  |  |
|  | 3-Chip image sensor (CCD or CMOS), Full HD. 4K to be priced separately |  |  |
|  | Compact, lightweight, easy to grip. |  |  |
|  | Resolution: a minimum of 1920*1080. |  |  |
|  | Focal Distance with zoom lens approx. 16-28 mm or better. |  |  |
|  | Gas sterilizable, Autoclavable is preferable. |  |  |
|  | Immersible in disinfections liquid. |  |  |
|  | Universal coupling mechanism to fit all standard endoscopes. |  |  |
|  | Camera cable approx. 3m. |  |  |
| B | FULL HD Camera Controller (console): |  |  |
|  | Full HD controller (4k to be priced separately) |  |  |
|  | Clear function buttons and indicators. |  |  |
|  | Hygienic. |  |  |
|  | Picture in picture mode. |  |  |
| C | Light Source: |  |  |
|  | LED Light source (light intensity should be equivalent to 300 Watt xenon lamp) |  |  |
|  | Working hour ≥ 10,000 hours. |  |  |
|  | Automatic and manual adjustment of light intensity. |  |  |
|  | Front panel hygienic with convenient indicators for ease of operation. |  |  |
|  | Flexible fibre optic cable preferably autoclavable and not less than 230 mm. |  |  |
|  | Provides optimal light transmission and long durability. |  |  |
| D | Suction/ Irrigation system: |  |  |
|  | Roller pump for irrigation (Pressure: 15-150 mmHg or better). |  |  |
|  | Flow Rate Adjustable: 0.1- 0.5 l/min or better. |  |  |
|  | Adjustable vacuum pressure. |  |  |
|  | Complete hysteroscopy set for irrigation system. |  |  |
| E | LCD / LED medical grade Monitor: |  |  |
|  | Full HD with a Resolution: minimum of 1920*1080, 4K to be priced separately |  |  |
|  | Medical grade type that is supplied by the same manufacturer. |  |  |
|  | A minimum of 26 inch. |  |  |
|  | To be compatible with the camera. |  |  |
| F | Mobile video cart: To suit the above systems and consists of the following: |  |  |
|  | 4 antistatic, anti-scratch castors. And four shelves |  |  |
|  | Locking brakes. |  |  |
|  | Cable channel with shockproof plugs. |  |  |
|  | Adjustable or fixed shelves. |  |  |
|  | Drawer. |  |  |
|  | All accessories (connect cable, plugs, etc.) needed for the offered system to be included. |  |  |
| G | High frequency electrosurgical generator |  |  |
|  | Controls to include mono-polar and bipolar modes, pure cut, blend cut, coagulate, and spray coagulation. |  |  |
|  | Bipolar Resection mode compatible with the bipolar resectoscope. |  |  |
|  | Automatic Stable output control to maintain selected power level into a wide range of different tissue types. |  |  |
|  | Neutral patient plate monitoring system. |  |  |
|  | Maximum CUT output power: 300-400 watt. |  |  |
|  | Digital display of power level and error codes. |  |  |
|  | The following accessories must be quoted separately: |  |  |
|  | Patient plate, adult, reusable Qty. (1) |  |  |
|  | Reusable patient plate cable. |  |  |
|  | Bipolar and monopolar cables. |  |  |
|  | Footswitch, twin type, spill proof. |  |  |
|  | All types of cutting and coagulating handles, tips and forceps for mono-polar and bipolar modes for endoscopic surgeries must be quoted separately. |  |  |
|  | Handle tips, forceps, cables and handles must be autoclavable. |  |  |
|  | Safety type: CF defibrillation proof. |  |  |
|  | Set of disposable adult Patient plate to be priced separately. |  |  |
| H | Medical Grade DVD Recorder Medical Grade DVD Recorder. Read & write functions. Internal Hard Drive capacity of 250 GB minimum. External USB Storage. |  |  |
| I | CO2 Insufflator: |  |  |
|  | Sealed hygienic front touch panel. |  |  |
|  | Display of settings. |  |  |
|  | Maximum flow rate From 40- 50 L/MIN. |  |  |
|  | Display of gas consumption, flow rate and intra-abdominal pressure |  |  |
|  | Gas heating to body temperature (at 37 C). |  |  |
|  | Audible and visual alarms. |  |  |
|  | Filter on the CO2 output to prevent cross-contamination. |  |  |
|  | Smoke evacuator that is either integrated or quoted as a separate unit. |  |  |

| Item 2 | Gynaecology Laparoscopic Surgical Set | Qty. (1 set) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |
| --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |
| FDA clearance OR CE Mark |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Needle holder, Ø 5.0 mm, jaw insert, straight Qty. (3) |  |  |
| 2 | Grasping & extraction forceps, Ø 5.0 mm, “Myoma”, single-action Qty. (3) |  |  |
| 3 | Grasping & extraction forceps, Ø 5.0 mm, “pike-mouth”, double-action Qty. (3) |  |  |
| 4 | Grasping & dissecting forceps, Ø 5.0 mm, “fenestrated”, single-action Qty. (3) |  |  |
| 5 | Grasping & dissecting forceps, Ø 5.0 mm, “Dolphin”, double-action Qty. (3) |  |  |
| 6 | Grasping & dissecting forceps, Ø 5.0 mm, right-angled, double-action Qty. (3) |  |  |
| 7 | Grasping & dissecting forceps, Ø 5.0 mm, long jaw, double-action Qty. (3) |  |  |
| 8 | Scissors “Metzenbaum”, curved, double-action Qty. (3) |  |  |
| 9 | Scissors, slimline, straight, double-action Qty. (3) |  |  |
| 10 | Laparoscopic telescope, 10 mm , 30° Qty. (1) |  |  |
| 11 | Laparoscopic telescope, 10 mm , 0° Qty. (1) |  |  |
| 12 | Laparoscopic telescope, 5 mm , 30° Qty. (1) |  |  |
| 13 | Laparoscopic telescope, 5 mm , 0° Qty. (1) |  |  |
| 14 | Hook electrode, monopolar, 5 mm Qty. (3) |  |  |
| 15 | Grasping & dissecting forceps, bipolar, “Maryland Dissector”, single-action Qty. (3) |  |  |
| 16 | Dissecting forceps, bipolar, non-fenestrated Qty. (3) |  |  |
| 17 | Biopsy forceps, 10 mm Qty. (3) |  |  |
| 18 | Crocodile forceps, 10 mm Qty. (3) |  |  |
| 19 | Any other instruments recommended by the manufacturer to be priced separately |  |  |

| Item 3 | Hysteroscopic Diagnostic Set | Qty. (1 set) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |
| --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |
| FDA clearance OR CE Mark |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Hysteroscopic telescope 30°, diameter ~ 2.9 mm, length ~ 30 cm Qty. (1) |  |  |
| 2 | Inner sheath diameter ~ 3.8 mm Qty. (1) |  |  |
| 3 | Outer sheath diameter ~ 4.5 mm Qty. (1) |  |  |
| 4 | Grasping forceps, “alligator jaws” Qty. (3) |  |  |
| 5 | Micro scissors Qty. (3) |  |  |
| 6 | Hooked scissors Qty. (3) |  |  |
| 7 | Any other instruments recommended by the manufacturer to be priced separately |  |  |

| Item 4 | Hysteroscopic Operative Set | Qty. (1 set) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |
| --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |
| FDA clearance OR CE Mark |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Resectoscope, slim ~ 7 mm, rotatable with working element, inner and outer sheathes Qty. (1) |  |  |
| 2 | Cutting electrode, bipolar, sterile Qty. (3) |  |  |
| 3 | Roller electrode, monopolar, sterile Qty. (3) |  |  |
| 4 | Hook electrode, monopolar, sterile Qty. (3) |  |  |
| 5 | Any other instruments recommended by the manufacturer to be priced separately |  |  |

SPECIAL TERMS

- Offers not complying with any of the special terms or the technical specifications shall be considered non-conforming with tender requirements.

- Any vendor providing FORGED documents shall be disqualified from the current tender and banned from participating in any future RMS tenders.

- All equipment must be the most recently released model/version which is equal to or higher than the range of the specifications of the required system (low, mid or high) and equal to or higher than the level of technology and required options mentioned in  the technical specifications.

- Required certificates:

- For equipment of US origin, a copy of a certificate of FDA approval & the relevant 510K clearance for selling to US healthcare facilities for the offered model must be submitted with the technical offer.

- For equipment of other origins, a copy of either a CE certificate with the relevant CE number (MDD)/TÜV/BSI/UL OR a certificate of FDA approval & the relevant 510K clearance for selling to US healthcare facilities for the offered model must be submitted with the technical offer.

- Only for class Ι medical equipment, submission of a copy of Declaration of Conformity certificate (MDD) for the offered model shall be accepted.

- With each offer, bidders must provide a formally endorsed document issued by the manufacturer stating that the bidder is the sole certified agent / distributor for the offered item.

- In all of the above cases (except 2.4) certificates must be formally endorsed by JFDA.

- Country of origin:

  - The country of origin of the main part (s) of the system must be one of the following:

USA, Canada, Japan, UK, Sweden, Finland, Denmark, Switzerland, Belgium, Germany, France, Netherlands, Spain, Norway, Italy, Ireland, Austria, New Zealand, Australia &Czech Republic.

  - Accessories and consumables may be manufactured in other countries and/or by different manufacturers.

  - All offered items must be approved for sale in the same country of origin. An original and officially endorsed free-sale certificate from an authorised body must be included in the offer.

  - Vendors must specify the origin of all offered items and accessories in the technical offer.

  - Except for equipment mentioned in (3.6) below, equipment manufactured by reputable companies based in any of the countries mentioned in (3.1) will be taken into consideration regardless of the manufacturing site only::

  - If they are approved for sale in the same county of origin (an original and officially endorsed free-sale certificate from an authorised body must be included in the offer.

OR

  - If they are approved for sale in at least three of the countries mentioned in (3.1) (an original and officially endorsed free-sale certificate from an authorised body in those countries must be included in the offer).

  - For X-ray based equipment, MRI, and nuclear medicine systems, the following parts must be manufactured in one of the countries mentioned in 3.1 above:

    - X-ray tubes

    - X-ray generators

    - Flat panel detectors

    - Gantries (including detectors)

    - Image intensifiers

    - MRI magnets

    - Gamma camera heads

  - Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first third of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period shall be rejected.

- Warranty:

  - Offers must include a full warranty including spare parts and labour for a period of a minimum of 24 months from the date of installation.

  - If at any time during the warranty period the item becomes inoperative due to a technical fault the item must then be repaired by the supplier /local agent within a period of fourteen days from written notification, otherwise the supplier must replace the item with a new identical functioning one and will endure a penalty determined by the Royal Medical Services for each day of the downtime of the system. In case the item was replaced by a new one, the warranty period mentioned in (4.i) above will start from the installation and commissioning date of the new item.

- One set of operation manual(s) and one set of service manual(s) including schematics and a spare-part list must be delivered with each unit, CD/DVD is acceptable. For large tenders, a certain agreed percentage of manuals per item may be agreed upon.

- Where applicable, pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include removal of old system(s), any civil work, electrical work or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new system(s).

- Power requirements: where applicable either single phase 220V, 50Hz or 3-phase 380V. Systems with external transformers are considered conforming only if clearly stated in the technical specifications.

- Technical offers must include clear original technical brochures/catalogues for all offered items.

- Offers must include fully detailed technical offers and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s), full technical description/specifications and any accessories or options included in the offer.

- Compliance sheets must be as per the tabular format of the technical specifications in the tender documents, listing the required specifications on one column and a  Yes  or NO  response to each point in the adjacent column, with reference to page and line numbers in the relevant technical brochure. Offers not complying with this term shall be rejected.

- Accessories and consumables:

  - Any accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately.

  - Technical offers must include a priced list for accessories and consumables as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) with prices fixed for a period of five years from the date of installation and commissioning with a maximum annual increase of 2%, any essential item not listed will be considered free of charge.

  - Accessories and consumables must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Where applicable, a start-up kit of accessories and consumable items must be provided with each system on a free-of-charge basis.

- Spare Parts:

  - Technical offers must include a comprehensive and priced spare parts list as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) valid for a minimum period of five years with a maximum annual increase of 2%, commencing at the end date of the warranty period, any essential item not listed will be considered free of charge.

  - Spare parts must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

- Spare parts, consumables and accessories availability must be guaranteed for a minimum period of ten years starting from the date of installation and commissioning.

- Tender Awards:

  - For the final list of offers having a chance of winning the award, the awarding process shall be based on the accumulative value of both the offered item and its’ running cost (Total Cost of Ownership) over a period of seven years from the date of installation and commissioning. Only offers with the lowest total cost of ownership over a period of seven years from the date of installation and commissioning shall qualify for the award.

  - Running cost includes the value of consumables, accessories needed to operate the system over the same period as well as the cost of any service contract (where applicable).

- For PC/Laptop based systems:

  - Complete restoration medium (CD/DVD/etc.) of the operating system and the application software must be supplied.

  - Where locally supplied computers or laptops are offered, only computers/laptops from Apple, hp/Compaq, Lenovo, Dell, fujtisu or Toshiba will be accepted, offered models must be the latest available version upon delivery.

  - Where locally supplied printers are offered only the following types and brands are accepted: HP, SAMSUNG, OKI, CANON, EPSON.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores, installation, pre-installation (if needed), training, commissioning, warranty and bringing the equipment into service.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- DRMS has the right to increase or decrease the awarded quantities by a percentage not exceeding 30% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- The supplier must furnish DRMS with a guarantee stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelvemonths from the date of final acceptance of the equipment by DRMS.

- Training: onsite user and service training

Special Terms for Surgical Instruments

- Offers not complying with any of the special terms or the technical specifications  will be considered as non-conforming.

- Any vendor providing FORGED documents will be disqualified from the current tender and any future RMS tenders or purchase orders.

- Bidder must provide a copy of the certificates of the following standards formally endorsed by the chamber of commerce of the manufacturing origin with the technical offer.

- EN ISO 13485:2012

(Standard to measure the quality of medical equipment, medical instruments and medical technology)

- ISO 9001:2008

(Specifies requirements for a quality management system)

- ISO 7153-1:1991

(This second edition cancels and replaces the first edition (ISO 7153-1:1983): it has been extended to include dental instruments)

- ISO 7151:1988

## (Specification of basic requirements for as well physical characteristics as workmanship, of the steel grades used and heat treatment of component parts, excluding rivets, screws)

- ISO 7740:1985

(Lays down the dimensions of two sizes of fitting features for detachable scalpel blades and the handles with which they are used)

- ISO 7741:1986

(This standard deals with materials, heat treatment and hardness of component parts, corrosion resistance, workmanship and cutting ability of scissors and shears used in the surgery and defines the test methods)

- ISO 13402:1995

(Describes test methods to determine the resistance of stainless steel surgical and dental hand instruments against autoclaving, corrosion and thermal exposure

- EN ISO 14001: 2004

ISO 14001 sets out the criteria for an Environmental Management System (EMS)

- 2007/47/EC

(Is intended to harmonize the laws relating to medical devices within the European Union)

- Bidder must submit samples with the technical offer from each brand that was submitted in the offer to be assessed and /or evaluated by RMS.

- Offers which do not include such samples will be considered non-conforming.

- Offers which fail the evaluation/assessment process will be excluded from the tender.

- Samples will be returned to the Bidder at the end of the evaluation/assessment process except for samples from Bidders winning the final award where they will be the property of RMS and will be kept as a reference for the receiving committee.

- Required certificates:

  - For instruments of USA origin, a copy of a certificate of FDA approval must be submitted with the technical offer.

  - For instruments of other origins, a copy of either a CE certificate (MDD)/TÜV OR a certificate of FDA approval must be submitted with the technical offer.

  - Where applicable, a copy of declaration of conformity certificate is accepted.

- Country of origin:

- The origin and the manufacturing plant for each set must be one of the following countries:     USA, UK, Sweden, Switzerland, Germany, France, Austria.

- All items should be engraved or etched with manufacturing origin, company logo and code number.

- Each instrument set will be awarded as a whole set.Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first third of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period will be rejected.

- The supplier must furnish DRMS with a guarantee letter stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelve months from the date of final acceptance of the equipment by DRMS which shall cover rusting, manufacturing defects, wear and tear and corrosion.

- Offers must include clear original catalogues for all offered items. Images and part numbers of offered items must be provided, highlighted and outlined clearly.

- Offers must include fully detailed information and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s),  full description/specifications.

- Compliance sheets must be as per the format provided with the tender documents for the specifications of the tender listing the required specification on one side and a Yes or NO response to each point on the other side with reference to page and line numbers in the specific brochure.

Offers not complying with this term will be rejected.

- Any accessories and consumable items necessary for the proper operation of surgical sets must be included in the offer. Bidder will bear any responsibility resulting from the awarding of insufficient items or sub-items.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores and warranty.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- DRMS has the right to increase or decrease the awarded quantities by a percentage not exceeding 30% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- For offers submitted in Jordanian dinars, payment will be either by wire transfer or by cheque after final acceptance of goods. Any other way of payment will be rejected.