| Annex 2 Anesthesia & General Common OR & Wards |  |  |
| --- | --- | --- |
| Item 1 | Operation Table | Qty. (18) |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| A | Operating Tables: |  |  |
| 1 | General operating table. It shall meet or exceed the following components/features: |  |  |
| 2 | Electromechanical or electro hydraulic mobile operating table column. |  |  |
| 3 | Radiolucent table-top: capability to be used with C-Arm, no metal bars in the middle of the table top for clear image |  |  |
| 4 | It shall have integrated, maintenance-free, rechargeable batteries with integrated battery charger. |  |  |
| 5 | 4-section operating table-top for general surgeries including the following: - Head section, removable - back section - Seat section - Leg section, pair, removable, with straps |  |  |
| 6 | Full function, corded hand control (wireless remote control to be offered as an option), the following functions must be included: - Height control (up & down) - Trendelenburg and reverse Trendelenburg - Tilt and back tilt - Flex & reflex - Zero position (return to level) - Leg movement |  |  |
| 7 | Override touch control panel with all remote control basic functions to be included and located on the table column for easy access. |  |  |
| 8 | Safe working load without any restrictions on all table 4 sections ≥ 300 kg |  |  |
| 9 | Approximate total length of the operating table-top with all of its sections: 2000 mm approximately. |  |  |
| 10 | Approximate table-top width (without side rails): 500 - 600 mm |  |  |
| 11 | Height (without mattress): ≤ 60 to ≥ 1000 mm. |  |  |
| 12 | Trendelenburg / reverse Trendelenburg: 30° or better |  |  |
| 13 | Lateral tilt (left / right): 20 ° or better |  |  |
| 14 | Back plate (up / down): +80° / -40° or better |  |  |
| 15 | Leg plate (up / down): +30°/ -100° with abduction movement |  |  |
| 16 | Motorized Longitudinal shift of at least 400 mm and should be controlled from the remote control |  |  |
| 17 | Table orientation: to work on normal and reverse orientations |  |  |
| 18 | Heavy duty stainless steel construction of table frame, side-rails and column casing or heavy duty aluminium alloy |  |  |
| 19 | European side rail |  |  |
| 20 | It shall have stable base design with shock resistant base cover |  |  |
| 21 | Four double antistatic swivel castors of at least 8 cm diameter with central braking system. |  |  |
| 22 | It shall have the possibility to set up multiple patient positions for general procedures. |  |  |
| 23 | Mattresses (complete set: head, upper back, lower back, pelvis, leg sections, arm boards) to be with at least the following specifications: - Thickness at least 80 mm - Anti-microbial - Solution resistant - Seamless (cover should be full-welded) - Waterproof |  |  |
| 24 | All general surgery options and accessories shall be included in the offer and priced separately |  |  |
| B | General Accessories: The following accessories and parts must be included in the offer with all needed radial clamps in addition |  |  |
|  | - Anaesthesia screen Qty. (16) |  |  |
|  | - Infusion pole Qty. (16) |  |  |
|  | - Body strap, multi size Qty. (16 set) |  |  |
|  | - Pair of arm boards, height adjustable, with straps Qty. (16 Pairs) |  |  |
|  | - Pair of knee crutches (Goepel design), can be mounted on the table without detaching the leg section Qty. (7 Pairs) |  |  |
|  | - Pair of lithotomy stirrups Qty. (2 Pairs) |  |  |
|  | - Pair of lateral support Qty. (11 Pairs) |  |  |
|  | - Pair of back support Qty. (9 Pairs) |  |  |
|  | - Pair of shoulder supports Qty. (9 Pairs) |  |  |
|  | - Gel pads, quote separately different options and sizes |  |  |
|  | - Headrest for prone position with all required mounting requirements Qty. (2) |  |  |
|  | - Accessories stand trolley Qty. (16) |  |  |
| C | Gynaecology Accessories: Soiled bowl Qty. (2) |  |  |
| D | Maxillofacial Accessories: Headrest horseshoe-shaped, one part, with all required mounting requirements Qty. (1) |  |  |
| E | ENT Accessories: Headrest horseshoe-shaped, two pats, with all required mounting requirements Qty. (2) |  |  |
| F | Vascular Accessories: Tabletop Extension: Qty. (1) - Carbon fibre, 360˚ radiolucent - Carbon length ≥ 110 cm - Width ≥ 50 cm - Pad to be included - Weight capacity ≥ 130 kg - All mounting accessories should be included |  |  |
| G | Neurosurgery/ Spine Accessories (priced separately with all mounting accessories): |  |  |
|  | G-1) Spine Extension System, Carbon fibre: Qty. (1) - A spine extension system to turn operating table into a specialty spine table that supports cervical, thoracic, and lumbar procedures with easy precise adjustment of the head and cervical spine - System should provide excellent C-Arm and O-Arm access - Spine System should provide intra-operative lumbar flex and repositioning of the patient - Small storage footprint - Radiolucent - Lateral tilt (clockwise and counter clockwise) ≥ 20° - Attachment to surgical table to be included - Patient weight capacity ≥ 200 kg - Height adjustable: 65 cm – 95 cm approximately - Mounted on sturdy column and base - Castors should be included - Tabletop Dimensions (L x W): ≈ 200 cm x 40 cm (in zero position) - It shall include at least the following with all needed mounting accessories on the extension (priced separately): |  |  |
|  | - All needed chest and hip supports with pads Qty. (1 set) |  |  |
|  | - Sling support to support patient legs Qty. (1) |  |  |
|  | - Supine tabletop set with pads for spine extension Qty. (1 set) |  |  |
|  | - Medical positioning pillows for prone position Qty. (1 Set) |  |  |
|  | - Prone protective head support Qty. (1) |  |  |
|  | - Pair of arm boards, height adjustable, with straps with all needed mounting accessories on side rail Qty. (1 Pair) |  |  |
|  | G-2) Wilson frame with pad protection and trolley Qty. (1) |  |  |
|  | G-3) Tabletop Extension: Qty. (1) - Carbon fibre, 360˚ radiolucent - Carbon length ≥ 110 cm - Width ≥ 50 cm - Pad to be included - Weight capacity ≥ 130 kg |  |  |
|  | G-4) Head positioning system for cervical traction (bidders are required to quote both types radiolucent carbon fiber and aluminum) , all adapters needed to be mounted on both table and spine extension should be included Qty. (1) |  |  |
|  | G-5) Patient Transfer Board (Sliding Board) 180 cm x 50 cm approximately Qty. (1) |  |  |
| H | Neurosurgery/ Cranial Accessories (priced separately): |  |  |
|  | H-1) Skull holder (bidders are required to quote both types radiolucent carbon fiber and aluminum) with reusable skull pins for adults and children Qty. (1) |  |  |
|  | H-2) Headrest, horseshoe shaped with gel pad, two parts Qty. (1) |  |  |
| I | Any other needed accessory needed for cranial and spine surgeries should be quoted separately |  |  |
| 25 | Abroad service training for at least 4 biomedical engineers/ technicians as per special terms |  |  |

| Item 2 | Conventional Electrosurgical Unit | Qty. (16) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well-known brand |  |  |
| 2 | Application: should be suitable for wide range of major surgeries (At least General, Orthopaedic, Laparoscopic, Caesarean, Spine, Cranial, Gynaecology, Vascular & ENT surgeries) |  |  |
| 3 | Controls to include mono-polar and bipolar modes, pure cut, blend cut, coagulate, spray coagulation |  |  |
| 4 | Automatic Stable output control to maintain selected power level into a wide range of different tissue types |  |  |
| 5 | Neutral patient plate monitoring system with indicator |  |  |
| 6 | Maximum Cut output power: 300-400 watt |  |  |
| 7 | Maximum Coagulation output power 80 watt or better |  |  |
| 8 | Leakage current ≤ 20 mA |  |  |
| 9 | Operating panel with touch buttons to select modes and adjust power levels. |  |  |
| 10 | Digital touch-screen LCD display of power level and error codes |  |  |
| 11 | Monopolar handle connection: International 3-pin plug type |  |  |
| 12 | The following accessories must be quoted separately: |  |  |
|  | - Patient plate, adult, split type, disposable |  |  |
|  | - Patient plate, paediatric, split type, disposable |  |  |
|  | - Patient plate, adult, reusable |  |  |
|  | - Patient plate, paediatric, reusable |  |  |
|  | - Reusable patient plate cable |  |  |
|  | - Bipolar forceps & cable |  |  |
|  | - Monopolar pencil, reusable |  |  |
|  | - Monopolar handle tips (Ball, loop, needle, spatula, curved spatula), all should be quoted separately |  |  |
|  | - Hook electrode with cable |  |  |
|  | - Footswitch, twin type, spill proof |  |  |
| 13 | All types of cutting and coagulating tips and forceps for mono-polar and bipolar modes for major surgeries (mentioned in 2) must be quoted separately. |  |  |
| 14 | Handle tips, forceps, cables must be autoclavable. |  |  |
| 15 | Control through push buttons on handle, and through foot switch. |  |  |
| 16 | Audible and visual alarm with self-test capability. |  |  |
| 17 | Small in size and should be able to be mounted on surgical pendant shelf without causing any interferences and signal distortion to nearby equipment |  |  |
| 18 | Safety type : CF defibrillation proof |  |  |
| 19 | Any options to be quoted separately |  |  |
| 20 | Offer must include one electrical safety analyser (priced separately): - From a reputable well known company - Low weight portable type with carrying handle - All standard accessories needed to run the analyser should be included |  |  |
| 21 | Abroad service training for 2 biomedical engineers/ technicians as per special terms |  |  |

| Item 3 | Transport Ventilator, Adult, Paediatric & Neonatal | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well known manufacturer |  |  |
| 2 | General Features |  |  |
|  | - Application: Adult / Paediatric/transport ventilator |  |  |
|  | - Microprocessor controlled |  |  |
|  | - Construction: Spill proof, Robust and corrosion resistant |  |  |
|  | - Self-test capability |  |  |
|  | - Trolley: - Stable design to prevent tipping over - Medical grade castors - Fitted with oxygen cylinder tray |  |  |
|  | - Open system for patient circuits |  |  |
|  | - Built in durable low noise turbine |  |  |
|  | - Ability to operate on high pressure sources |  |  |
|  | - Unit must be easy to dismount from the trolley and carry around, or place on patient stretcher. |  |  |
|  | - Gas hoses (Oxygen) - Original make - Colour coded - Special hose for connection to oxygen cylinder |  |  |
|  | - Display: Min 8 inches, TFT full colour, touch-screen |  |  |
|  | - An upper handle to easily transport the ventilator. |  |  |
| 3 | Ventilation Modes and Functions |  |  |
|  | - Assist / Control mode: volume and pressure |  |  |
|  | - SIMV |  |  |
|  | - Spontaneous CPAP |  |  |
|  | - Pressure support ventilation |  |  |
|  | - Additional modes: please specify |  |  |
| 4 | Controls and Settings |  |  |
|  | - Tidal Volume (ml): minimum range for adult/ pediatric: (50-2000) ml minimum range for neonates: (5-200) ml |  |  |
|  | - Inspiratory flow (l/min) minimum range : 10-80 L/min |  |  |
|  | - Respiration rate (bpm): (5-50) or better |  |  |
|  | - Inspiratory time: specify |  |  |
|  | - Pressure control: min range 5-50 cm H2O |  |  |
|  | - PEEP: min range 0-20 cm H2O |  |  |
| 5 | Control Panel for Parameters |  |  |
|  | - Inspiratory time |  |  |
|  | - I:E ratio |  |  |
|  | - PEEP |  |  |
|  | - Respiration frequency / minute |  |  |
|  | - FiO2 |  |  |
|  | - tidal volume |  |  |
| 6 | Patient Alarms |  |  |
|  | Audible and visual user adjustable alarms for : |  |  |
|  | - Low / high inspiratory pressure |  |  |
|  | - Low / high FiO2 |  |  |
|  | - High/Low tidal volume |  |  |
|  | - Apnea time |  |  |
|  | - Breathing circuit disconnection |  |  |
| 7 | Equipment Alarms |  |  |
|  | - Low battery |  |  |
|  | - Gas supply loss |  |  |
|  | - Power failure |  |  |
|  | - Ventilator inoperative |  |  |
|  | - Continuous audible alarm for at least one minute in case of mains power failure even when the battery is depleted |  |  |
| 8 | Electric Requirements |  |  |
|  | - Mains: - Voltage: 220 V a.c. - Frequency: 50 Hz - Power cable: original British standard cable and plug - Ability to work on ambulance power outlet |  |  |
|  | - Battery: - Back-up time for settings and operation not less than 4 hours - Batteries should be easy replaceable - Battery type and specifications should be mentioned clearly, bidder should state clearly if the battery is available in the local market; if else it price of the battery should be mentioned in the technical offer and fixed as per special terms - Any required battery conditioner should be included |  |  |
| 9 | Standard Accessories List all standard accessories required for commissioning the system and indicate if they are included in the main unit price |  |  |
| 10 | Optional Accessories List all optional accessories and price them separately including oxygen cylinders |  |  |
| 11 | Abroad service training for 1 biomedical engineer/ technician as per special terms |  |  |

| Item 4 | Pneumatic Sequential Compression System for DVT | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Intermittent pneumatic compression system to prevent deep vein thrombosis (DVT) |  |  |
| 2 | LCD display |  |  |
| 3 | Light weight and small in size with carrying handle |  |  |
| 4 | Liquid proof design with at least IPX2 rating |  |  |
| 5 | Low noise operation |  |  |
| 6 | All standard accessories must be included (tubing and sleeves) and must be priced separately with all available sizes and types (leg and foot) |  |  |

| Item 5 | Tourniquet, Double Cuff, Electrical | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Dual Channel |  |  |
| 2 | Electrically operated with built in air compressor |  |  |
| 3 | Digital LCD display ≥ 6 inches of preset and actual pressure in both channels. |  |  |
| 4 | Built in digital timer with timer alarm after reaching the set alarm time |  |  |
| 5 | Built-in backup rechargeable battery: at least 2 hours |  |  |
| 6 | Adjustable pressure range: 100 to 500 mmHg or better |  |  |
| 7 | With leak test capability |  |  |
| 8 | Mounted on robust mobile stand and basket for cuffs and accessories |  |  |
| 9 | Cuffs should be reusable (up to at least 100 times), durable, of high quality and colour coded and should not be damaged by disinfecting with commonly used agents used in the operating theatres |  |  |
| 10 | All control unit output ports, cuffs and cuff tubing should have locking connectors. |  |  |
| 11 | All types and sizes of cuffs should be listed and priced separately |  |  |
| 12 | Other accessories and consumables to be listed and priced separately |  |  |

| Item 6 | Cabinet, Warming, Dual, Floor | Qty. (8) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Dual heating chambers for warming of linens, blankets, IV fluids |  |  |
| 2 | Free-standing type |  |  |
| 3 | Capacity ≥ 400 litres |  |  |
| 4 | It should be constructed from heavy duty stainless steel with stainless steel shelves and doors |  |  |
| 5 | Digital readout of temperature inside cabinet in Celsius |  |  |
| 6 | Uniform heat transfer in chambers |  |  |
| 7 | It should have heating range from approx. 37°C to approximately 70°C. |  |  |
| 8 | High & low temperature alarm |  |  |

| Item 7 | Diagnostic Set | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, compact, heavy duty construction & reliable, ergonomic design and easy to use |  |  |
| 2 | Power handle, Ophthalmoscope head and an Otoscope head with at least three different sizes of reusable specula. |  |  |
| 3 | A hard-shell well-padded protective case. |  |  |
| 4 | Power handle should be made from heavy duty, easy to clean, corrosion resistant material. |  |  |
| 5 | Batteries and light bulbs should be commonly available from several known manufacturers. |  |  |
| 6 | Kindly quote for regular and/or rechargeable battery types. |  |  |
| 7 | Fibre-optic cold light transmission |  |  |
| 8 | LED illumination type |  |  |
| 9 | Kindly quote for all sizes and types of specula / accessories. |  |  |
| 10 | Full display of lens values |  |  |
| 11 | Wall mounted option should be quoted |  |  |

| Item 8 | Stethoscope | Qty. (14 Adult) Qty. (2 Paediatric) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, compact, heavy duty construction & reliable, ergonomic design and easy to use |  |  |
| 2 | Dual sided stainless steel chest piece |  |  |
| 3 | Lightweight |  |  |
| 4 | Latex free Y-tube |  |  |
| 5 | High quality diaphragm material with non-chill rim for patient comfort |  |  |
| 6 | Comfortable replaceable earpieces. |  |  |
| 7 | - Qty. (14) of the total quantity for adults - Qty. (2) of the total quantity for paediatrics |  |  |

| Item 9 | Anaesthesia Machine with Ventilator & Patient Monitor, on Pendant | Qty. (18) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction & reliable from reputable well-known brand, it should be suitable for adults, paediatrics and neonates. |  |  |
| 2 | It should be capable of providing low-flow techniques to minimize gas and anesthetic agent consumption. Guidance system to set the fresh gas flow to ensure minimum required fresh gas sufficiency during low flow. |  |  |
| 3 | Gas Supply: - Central gas supply (Pipeline gas inlets) all with international colour code and pressure gauge according to NIST system and equipped with B.O.C probes for connection to central gas outlets (O2, N2O & Medical Air) - Cylinder gas supply: Ability to be connected to external cylinders (O2, N2O & Medical Air) mounted on the anaesthesia gas pendant, all hoses to be included |  |  |
| 4 | The following should be included with the anaesthesia unit: - Upper shelf - Working tray/ surface - At least 2 large drawers - 4 service wheels |  |  |
| 5 | Flow meters: Machine should provide electronic gas mixing with digital control for O2, N2O and Air |  |  |
| 6 | Hypoxia guard (O2/N2O safety control): The unit should be equipped with Integrated Ratio System to maintain at least 21 % O2 in Fresh Gas whenever the oxygen supply pressure is reduced below normal or when accidental opening of only N2O flow with O2 valve closed |  |  |
| 7 | Oxygen-flush range: up to 50 L/min |  |  |
| 8 | Breathing Module: - Fully autoclavable module to deliver fresh gas to the patient (materials in contact with exhaled patient gases are autoclavable), it should be easy removable and reinstalled without tools - Flow sensor: Reusable type - It should have Pressure Graduated APL Valve (adjustable up to at least 70 cmH2O), Inspiratory Valve, Expiratory Valve & Active Gas Scavenging Port - Breathing system warmer to minimize humidity in the breathing system |  |  |
| 9 | CO2 absorber with a reusable canister of approximately 0.7 kg or more, the canister should be easy removable and reinstalled without tools |  |  |
| 10 | Suction regulator with all required hoses and probes (BOC) |  |  |
| 11 | Common fresh gas outlet (auxiliary gas outlet): The unit should have common gas outlet for using open circuit |  |  |
| 12 | Anesthetic vaporizers (manufactured by the same manufacturing company of the anesthesia machine): Offer to include Selectatec type (or equivalent) maintenance free Isoflurane & Sevofluorane vaporizers (mounted at the same time), all keys and adapters should be included |  |  |
| 13 | Self-test: Machine should perform automatic self-test to at least the following: - System test - Circuit leak - Calibration of all sensors |  |  |
| 14 | Ventilator: |  |  |
|  | - Integrated Microprocessor controlled ventilator, Ventilator brand must be the same as the anaesthesia unit and manufactured by the same company |  |  |
|  | - Ventilation modes to be selectable and to include at least: - Manual/ Spontaneous Ventilation - VCV (Volume Controlled Ventilation) - PCV (Pressure Control Ventilation) - PSV (Pressure Support Ventilation) - SIMV (Synchronized Intermittent Mechanical Ventilation) |  |  |
|  | - Fully coloured touch-screen LCD display of not less than 15" size to monitor (numeric & graphics): - Real time FiO2, capnograph, Anesthetic agents Isoflurane/ Sevofluorane & N2O (if not shown on the patient monitor) - Tidal Volume, Minute Volume, Respiration Frequency, PEEP & Mean pressure - MAC: It should have a display of MAC (Minimum Alveolar Concentration) - Alarms |  |  |
|  | - Ventilators Technical Data: - Tidal Volume: 5-1500 ml or better - Inspiratory pressure (P_insp): 5 to 60 cmH2O or better - Pressure limitation (P_max): 12 to 80 cmH2O or better - Respiratory Rate: 4 to 100 breaths/ minute or better - Inspiratory/ expiratory ratio: 1:4 to 4:1 or better - Positive End Expiratory Pressure (PEEP): 4 to 20 cmH2O or better |  |  |
|  | - Piston or bellow type ventilator |  |  |
|  | - For bellow type ventilator, bellows must be not other than ascending type |  |  |
| 15 | Alarms, audible and visual alarms to include: - Power Failure - O2 inlet supply failure - Low/ High Tidal Volume - Low Inspired oxygen (FiO2) - Apnea - Low/ High airway pressure |  |  |
| 16 | Built in rechargeable battery: - Minimum of 45 min. operation when fully charged. - Battery price should be mentioned in the technical offer and fixed as per special terms - Battery should be replaced in free of charge basis during the warranty period - Any required battery conditioner should be included |  |  |
| 17 | Accessories - Reusable Closed-type patient circuits suitable for adults, paediatrics & Neonates - Gas Cylinders: Qty. (2) oxygen cylinders (size E) with all required hoses and regulators - All required mounting accessories on the anaesthesia pendant and should be compatible with the model purchased in modular rooms. |  |  |
| 18 | Physiological Patient Monitor: Configuration: modular type with a minimum of three slots for three plug-in modules that can be fitted simultaneously on a module rack/ slots. - LCD display ≥ 15 inch touch-screen - Multi parameter module to measure the following parameters: ECG, Respiration rate, NIBP, Body temperature, SPO2 (Masimo technology) & IBP - Additional IBP (outlet or separate module) to be quoted separately - Depth of Anesthesia Monitoring Module (to be quoted separately) Qty. (2) modules for the whole required monitors, all accessories should be included, disposables should be quoted separately with prices fixed as per special terms - Cardiac output modules (CO) to be quoted separately Qty. (2) modules for the whole required monitors, all accessories should be included, disposables should be quoted separately with prices fixed as per special terms - Neuromuscular Transmission Module (NMT): Qty. (1) module or stand alone equipment for each monitor, all accessories should be included, disposables should be quoted separately with prices fixed as per special terms - Complete with all accessories for all required parameters (adult & pediatric) - Built-in rechargeable battery minimum life 1 hour of continuous monitoring - Min 72 hours trend capacity - QT and QTc analysis - ST Histogram - Data such as waveforms, measurement numerics and alarms from the anaesthesia machine can be displayed on the monitor for complete clinical overview - Escalating level alarm that cannot be permanently silenced - It should be securely mounted on the anaesthesia machine |  |  |
| 19 | Full gas module (on each machine): either built-in or removable on the anaesthesia machine to measure O2 (measured by permanent paramagnetic sensor), CO2, N2O & anaesthetic agents (Isoflurane & Sevofluorane), all necessary accessories to be included. Water trap: - must be priced separately in the technical offer, price should be fixed as per special terms - lifetime should be indicated in the technical offer |  |  |
| 20 | Central Station Qty.(2) |  |  |
|  | - Bed licenses: 8 bed licenses |  |  |
|  | - The central station is to be located at the anaesthesiologist office. Bidders shall quote for a comprehensive offer including network and any required cabling, switches, etc….. |  |  |
|  | - Central station must be supplied by the same manufacturer of the anaesthesia machines, locally supplied central stations will be considered as non-conforming. |  |  |
|  | - The central station monitor shall display waveforms for all parameters in both anaesthesia machines and patient monitors |  |  |
|  | - Ability to display waveforms, numeric data, graphic displays, tabular displays, and calculations. |  |  |
|  | - Audio / visual alarms with standard alarm / acknowledge / reset control. |  |  |
|  | - High resolution TFT-LCD touch screen colour monitor minimum 19 inches |  |  |
|  | - Ability to store patient data for at least 96 hours. |  |  |
|  | - HL7 License (compatible with Hakeem hospital information system) |  |  |
|  | - All available options and accessories must be offered and priced separately |  |  |
| 21 | Gas Analyser: Offer must include one gas analyser (priced separately, one unit with the whole quantity) from well-known brand, compact unit with the ability to measure the following: - High & low pressure & Ultra low pressure - Flow & Ultra low flow - Tidal volume - O2, CO2, N2O measurements - Anaesthetic Agents (isoflurane & Sevofluorane) measurements - Analyser should be compatible with anaesthesia machines from different manufacturers - All accessories needed to run the analyser should be included |  |  |
| 22 | Abroad service training for 4 biomedical engineers/ technicians as per special terms |  |  |

| Item 10 | Cerebral/Somatic Oximeter Monitor | Qty. (3) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Application: The monitor should measure regional tissue oxygen saturation of blood (rSO2) in the brain and skeletal muscle tissues of the body |  |  |
| 2 | It may be used for cerebral oximetry, somatic oximetry or both simultaneously noninvasively and on real time |  |  |
| 3 | It can be used for adult, pediatric and neonatal patients in any clinical setting where the brain and body are at risk of reduced flow or no-flow ischemic states |  |  |
| 4 | At least 2 channels |  |  |
| 5 | Extra 2 channels to be priced separately |  |  |
| 6 | User friendly system |  |  |
| 7 | At least 24 hours trend memory |  |  |
| 8 | rSO2 range: 15 – 95 or better |  |  |
| 9 | Ability to display waveforms and numeric data |  |  |
| 10 | High resolution colored LCD display of size not less than 10 inches |  |  |
| 11 | Back-up built in rechargeable battery with at least 20 minutes capacity |  |  |
| 12 | Audio and visual alarms (low & high rSO2) |  |  |
| 13 | The system should be compatible with HAKEEM hospital information system; all licenses, software and hardware needed should be included |  |  |
| 14 | The monitor should have the capability to be mounted on pendants inside the operation theatre, all accessories should be included |  |  |
| 15 | All required accessories needed to run the system should be included |  |  |
| 16 | All required consumables including all sensor types should be listed and priced separately |  |  |
| 17 | The following options should be quoted separately if available including all needed software, licenses, hardware & consumables, prices should be fixed as per special terms: - Cardiac output monitoring - Depth of anesthesia monitoring - Hemoglobin monitoring |  |  |
| 18 | All other options to be listed and priced separately |  |  |

| Item 11 | Defibrillator/Monitor | Qty. (16) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well-known manufacturer, durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Dedicated for hospital use only |  |  |
| 3 | Easy to use defibrillator/monitor with both manual and AED defibrillation capabilities. |  |  |
| 4 | Biphasic waveform technology. |  |  |
| 5 | Arrhythmia ECG monitoring capability through 3, 5 ECG lead sets and through external paddles. |  |  |
| 6 | Synchronisation capability through front panel button |  |  |
| 7 | Lightweight unit of less than 8 kg with batteries and external paddles |  |  |
| 8 | Coloured LCD display of size not less than 5.5 inches and with at least the following: |  |  |
|  | - Error messages (system errors, disconnected ECG lead or pads etc.) |  |  |
|  | - Heart rate display |  |  |
|  | - Selected energy level display |  |  |
|  | - Synchronization indicator display |  |  |
|  | - Battery capacity status and AC power supply indicator |  |  |
| 9 | Energy selection: Through front panel rotary knob or touch buttons to select energy level between 2 Joules and 200 Joules (or higher) |  |  |
| 10 | Shock control: Through both front panel and paddles. |  |  |
| 11 | Charging time: less than 7 seconds to reach 200 Joules energy level using new fully charged batteries.` |  |  |
| 12 | Discharge Time: must be specified, preference will be given to shorter time |  |  |
| 13 | Patient Impedance Range: Should compensate on patient impedance of range not less than: 25 to 200 Ω |  |  |
| 14 | The unit must feature charging and charged indicator tones |  |  |
| 15 | QRS beeper with adjustable volume |  |  |
| 16 | Self-test defibrillator capability |  |  |
| 17 | Built in rechargeable batteries with the following specifications: |  |  |
|  | - Less than 5 hours charging time |  |  |
|  | - Capacity: at least 100 shocks at 200 Joules energy level |  |  |
|  | - Batteries should be easy accessible & replaceable without any tools |  |  |
|  | - Battery type and specifications should be mentioned clearly, bidder should state clearly if the battery is available in the local market; if else it price of the battery should be mentioned in the technical offer and fixed as per special terms |  |  |
|  | - Battery should be replaced in free of charge basis during the warranty period |  |  |
|  | - Any required battery conditioner should be included |  |  |
| 18 | Built in full annotation recorder with 25 mm-sec nominal speed. |  |  |
| 19 | The unit must feature audio and visual alarms |  |  |
| 20 | Fluids protection level: IPX 1 or better |  |  |
| 21 | Defibrillation protection proof (patient isolation) type CF to ECG cables and internal paddles |  |  |
| 22 | Defibrillation protection proof (patient isolation) type BF to external paddles and disposables pads. |  |  |
| 23 | The following must be included with each unit: |  |  |
|  | - 3 lead ECG cable Additional cable and trunk should be priced separately |  |  |
|  | - External adult paddles, adult plates slide off to expose paediatric electrode surface Qty. 1 set |  |  |
|  | - Paper roll Qty. 10 Paper rolls should be available in the local market |  |  |
| 24 | The following must be quoted and priced separately: |  |  |
|  | - 5 Lead ECG cable |  |  |
|  | - Multifunction electrode pads |  |  |
| 25 | All Accessories including ECG cables, disposable electrodes, external paddles, paper rolls and electrode pads should be priced separately; prices should be fixed for at least 5 years after warranty period |  |  |
| 26 | Any other accessories or options should be quoted and priced separately |  |  |
| 27 | Offer must include one Defibrillator Analyser (priced separately): - From well-known manufacturer - With multi load - Compatible with different defibrillator manufacturers - Complete with all standard accessories |  |  |
| 28 | Abroad service training for 2 biomedical engineers/ technicians as per special terms |  |  |

| Item 12 | Ultrasound Machine, Portable | Qty. (8) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | Anaesthesia dedicated ultrasound scanner to be used in operating theatres. The system must meet or transcend the following specifications: |  |  |
| A | # Application: Abdominal, Vascular, Transcranial, Musculoskeletal OB/GYN, Small parts, nerve application |  |  |
| B | ## Configuration |  |  |
|  | ## Portable design with original trolley |  |  |
|  | ## Four full swivel castors with brakes |  |  |
|  | ## High-resolution LCD: ≥ 12” |  |  |
|  | ## Storage for transducer cables |  |  |
|  | ## Minimum of 2 active transducer connector on trolley |  |  |
|  | ## User-adjustable presets |  |  |
|  | ## Adjustable monitor position |  |  |
| C | General Features: |  |  |
|  | - Grays-scale levels: 256 |  |  |
|  | - Power up time: ≤ 2 minutes |  |  |
|  | - Shutdown confirmation capability |  |  |
|  | - Self-diagnostic routine during power up |  |  |
|  | - Acquisition frame rate: ≥ 500 fps |  |  |
|  | - Simultaneous processing bandwidth capability: 2-12 MHz |  |  |
|  | - No. of digitally processed channels: ≥ 100000 |  |  |
|  | - Input dynamic range: : ≥ 160 dB |  |  |
|  | - Cine loop: ≥ 200 frames |  |  |
|  | - Automatic real-time Doppler trace capability including calculation and display of user-selected measurements |  |  |
|  | - Real-time Pan/Zoom |  |  |
|  | - Frozen image Pan/Zoom |  |  |
|  | - Automatic optimization of base line and scale |  |  |
|  | - Local HDD patient search through name and/or ID |  |  |
|  | - Full DICOM 3.0, RIS, HIS, PACS Support |  |  |
|  | - Raw data processing capability or equivalent |  |  |
| D | Image Storage: |  |  |
|  | - Capacity of stored images HDD, please specify. |  |  |
|  | - Removable Storage: CD/DVD, USB |  |  |
| E | Multi-frequency Transducers: |  |  |
|  | - Convex probe: 2-5 MHz |  |  |
|  | - Linear probe: 5-12 MHz |  |  |
| F | Imaging Modes: |  |  |
|  | - 2D |  |  |
|  | - B-mode |  |  |
|  | - Tissue harmonic imaging |  |  |
|  | - Colour Doppler imaging |  |  |
|  | - Zoom function with Movable zoom box. |  |  |
|  | - Speckle reduction ( Noise reduction imaging) |  |  |
|  | - Needle guide imaging software |  |  |
| G | Reporting and Management System |  |  |
|  | - The system shall be capable to perform patient report which should include; patient data, measurements, analyses, to allow embedding of images into patient reports. |  |  |
|  | - Thermal Black and White printer |  |  |
| H | All available standard & optional features, packages, & accessories must be listed and priced separately |  |  |
| I | Abroad service training for 1 biomedical engineer/ technician as per special terms |  |  |

| Item 13 | Volumetric Infusion Pump | Qty. (28) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well-known manufacturer, durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Performance data: |  |  |
|  | - Flow rate setting: 0.1-1200 ml/hr or better |  |  |
|  | - Volume (VTBI) setting 1-9999 ml or better |  |  |
|  | - Automatic and manual bolus administration that is adjustable in rate and volume |  |  |
|  | - KVO (keep vein open) rate: adjustable 0.1-10 ml/hr or better |  |  |
|  | - Prime (purge) function |  |  |
|  | - Accuracy ± 5% (or better) of displayed flow rate and volume |  |  |
|  | - Automatic pressure should be reduced when occlusion detected |  |  |
|  | - Air bubble detector |  |  |
|  | - The pump should be an open system (major set suppliers) while maintaining accuracy. |  |  |
| 3 | Pump Mechanism: either linear peristaltic or (Linear, dual stage, positive displacement, flow compensated ) |  |  |
| 4 | Microprocessor controlled unit with start-up self-test capability with an error code system |  |  |
| 5 | Clear and easy to read LCD display of the following: |  |  |
|  | - Flow rate |  |  |
|  | - Volume infused and volume to be infused |  |  |
|  | - Time remaining |  |  |
|  | - Battery capacity |  |  |
|  | - Pressure level |  |  |
|  | - Bolus delivery |  |  |
|  | - Alarms and error codes |  |  |
| 6 | Hygienic, flat, soft touch keys control panel with visible and clearly identified controls |  |  |
| 7 | Battery and mains AC power supply indicators |  |  |
| 8 | Built in maintenance free rechargeable battery that is easy accessible with capacity not less than 6 hours operation at 25 ml/h flow rate when fully charged. |  |  |
| 9 | Visual and audible alarms with adjustable alarm volume control, The following alarms should be included: |  |  |
|  | - Upstream and downstream occlusion. |  |  |
|  | - Malfunction alarms. |  |  |
|  | - Air bubble. |  |  |
|  | - Low battery alarm. |  |  |
|  | - Door open |  |  |
|  | - infusion complete dose |  |  |
| 10 | Easy and secured clamping mechanism for pole fixation |  |  |
| 11 | Light weight (≤ 3 kg) |  |  |
| 12 | To work on either internal or external power supply |  |  |
| 13 | A list of standard accessories for the offered model. |  |  |
| 14 | Please quote for any other optional accessories; IV pole, docking station, etc... |  |  |

| Item 14 | Syringe Pump, Anaesthesia | Qty. (34) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well-known manufacturer, durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Performance data: |  |  |
|  | - TCI (Target controlled infusion) mode for TIVA (Total intravenous anaesthesia). |  |  |
|  | - Standard mode |  |  |
|  | - Flow rate setting: 0.1-999 ml/hr or better |  |  |
|  | - Volume (VTBI) setting 0.1-999 ml or better |  |  |
|  | - Drug dose rate calculation function |  |  |
|  | - Automatic and manual bolus administration that is adjustable in rate and volume |  |  |
|  | - KVO (keep vein open) function |  |  |
|  | - Prime (purge) function |  |  |
|  | - Accuracy ± 2% (or better) of displayed flow rate and volume |  |  |
|  | - Adjustable limit for occlusion pressure. |  |  |
|  | - Automatic pressure should be reduced when occlusion detected |  |  |
|  | - To accept syringe sizes from 10 ml up to 50 ml (or better) from several manufacturers (major suppliers) while maintaining accuracy (open system). |  |  |
| 3 | Pump Mechanism: Syringe driver |  |  |
| 4 | Microprocessor controlled unit with start-up self-test capability with an error code system |  |  |
| 5 | Medication database: pre-programmed drug protocols |  |  |
| 6 | Clear and easy to read LCD display of the following: |  |  |
|  | - Flow rate |  |  |
|  | - Volume infused and volume to be infused |  |  |
|  | - Time remaining |  |  |
|  | - Battery capacity |  |  |
|  | - Pressure level |  |  |
|  | - Bolus delivery |  |  |
|  | - Medication name and settings |  |  |
|  | - Alarms and error codes |  |  |
| 7 | Hygienic, flat, soft touch keys control panel with visible and clearly identified controls |  |  |
| 8 | Battery and mains AC power supply indicators |  |  |
| 9 | Built in maintenance free rechargeable battery that is easy accessible with capacity not less than 6 hours operation at 5 ml/h flow rate when fully charged. |  |  |
| 10 | Visual and audible alarms with adjustable alarm volume control, The following alarms should be included: |  |  |
|  | - High pressure occlusion. |  |  |
|  | - Syringe unlocked |  |  |
|  | - Malfunction alarms. |  |  |
|  | - Low battery alarm. |  |  |
|  | - Empty and near empty syringe |  |  |
| 11 | Easy and secured clamping mechanism for pole fixation |  |  |
| 12 | Light weight (≤ 3 kg) |  |  |
| 13 | To work on either internal or external power supply |  |  |
| 14 | A list of standard accessories for the offered model. |  |  |

| Item 15 | Blood Warmer | Qty. (16) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well-known manufacturer, durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | It shall be a microprocessor-controlled blood warmer of all infusions and transfusions in routine and massive application. |  |  |
| 3 | Temperature setting range 37° C - 41° C in ≤ 0.5 increments. |  |  |
| 4 | Automatic over heat safety switch off |  |  |
| 5 | Audible/visual alarms and indicators for high temperatures |  |  |
| 6 | Splash proof panel |  |  |
| 7 | Light weight ≤ 3 kg |  |  |
| 8 | Heated sleeve to avoid blood cooling from warmer to patient |  |  |
| 9 | Universal fixation for fast and safe fixation on IV poles |  |  |
| 10 | Direct AC 220 volts power supply without external power converters. |  |  |
| 11 | The unit should operate on disposables from different manufacturers |  |  |
| 12 | 20 original sets of all required consumables (including PVC blood line, air bubble trap and any other required consumables) for each system must be listed in a clear and itemized manner and priced separately. |  |  |

| Item 16 | Patient Warmer, Air Type | Qty. (22) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well-known manufacturer, durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Main purpose: whole body patient warming therapy to maintain normal body temperature during surgical operation and recovery. |  |  |
| 3 | Microprocessor controlled |  |  |
| 4 | Forced air convective warming technology that delivers gently moving warm air |  |  |
| 5 | Soft touch key control panel or touch LCD display to select air temperature. |  |  |
| 6 | At least three temperature selection: 32° C, 38° C, 43° C (± 3° C) |  |  |
| 7 | Fast warming operation |  |  |
| 8 | Safety temperature cut off to protect the heaters |  |  |
| 9 | Robust, heavy duty construction and compact small foot print design |  |  |
| 10 | Mobile on at least 4 antistatic castors |  |  |
| 11 | Smooth quiet operation of less than 55 db noise level |  |  |
| 12 | Highest level of safety for both patient and machine with alarm indicators. |  |  |
| 13 | High and low temperature alarms. |  |  |
| 14 | High efficiency ≤ 0.3 µm air filter |  |  |
| 15 | Light weight |  |  |
| 16 | All blanket sizes to be priced separately |  |  |
| 17 | Prices of blankets must be fixed as per special terms |  |  |
| 18 | The system must be complete with all standard accessories needed to be operated. |  |  |

| Item 17 | Suction Machine | Qty. (24) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Mobile mounted on antistatic castors (at least 4) with sturdy stable base that resists tipping over |  |  |
| 3 | On/Off switch with light indicator. |  |  |
| 4 | Maximum flow rate not less than 50 l/min. |  |  |
| 5 | Easy control of vacuum level through a knob or a switch on the control panel. |  |  |
| 6 | Maximum vacuum not less than 675 mmHg. |  |  |
| 7 | Display of vacuum through a gauge that is easy to read with high visibility mmHg and Kpa graduation. |  |  |
| 8 | 2 original reusable and autoclavable jars to be included with the unit, with lids that provide secure seal and float-type cut off valve to prevent overflow (All Jar sizes to be offered). |  |  |
| 9 | Bacteria filter to minimize cross infection (available from different manufacturers) |  |  |
| 10 | Silicon transparent tubes. |  |  |
| 11 | Change-over valve to switch suction between jars. |  |  |
| 12 | Splash proof design. |  |  |
| 13 | Oil-free maintenance free unit |  |  |
| 14 | Quiet pump operation (≤45 dB) |  |  |
| 15 | Any other accessories must be listed and priced separately. |  |  |
| 16 | Two extra original Jars are to be provided with the unit and priced separately |  |  |

| Item 18 | Resuscitation Bag (different sizes) | Qty. (65 Adult) (26 Paediatric) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, comfortable and reliable single layer bag |  |  |
| 2 | Complete with reservoir bag, pressure relief valve, reservoir one way valve and air cushioned face mask |  |  |
| 3 | All parts (mentioned above) should be easily assembled/ disassembled, easy to clean and disinfect and autoclavable at 134° C |  |  |
| 4 | All parts mentioned above should be made silicon made and latex free. |  |  |
| 5 | 100 % O2 concentration (supply delivery) with O2 accumulator. |  |  |
| 6 | - Adult size Qty. (65) - Paediatric Size Qty. (26) |  |  |
| 7 | All accessories to be mentioned and priced separately. |  |  |

| Item 19 | Laryngoscope Set, Adult, Pediatric & Neonate | Qty. (20 Adult) Qty. (5 Pediatric) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Light weight. |  |  |
| 3 | LED type. |  |  |
| 4 | Fibre optic plate illumination type. |  |  |
| 5 | To be made from heavy-duty stainless steel (Handle & Blades) |  |  |
| 6 | Standard size, non-rechargeable battery operation. |  |  |
| 7 | Durable case with sufficient padding. |  |  |
| 8 | - Qty. (20) of the total quantity with 2 curved blades each, sizes: 3, 4 (for adult patients) Size 5 to be priced separately - Qty. (5) of the total quantity with 3 blades each, sizes: 1, 2 straight & size 2 curved (for paediatric patients) |  |  |

| Item 20 | Video Assisted Laryngoscope | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| A | Video laryngoscope unit, for the use of difficult intubation for adults & paediatrics |  |  |
| 1 | Monitor: |  |  |
|  | - Size ≥ 6 inches |  |  |
|  | - High resolution LCD type |  |  |
| 2 | Rigid blades: |  |  |
|  | - With built in video camera: CMOS technology that provides superior video image quality |  |  |
|  | - Built in LED light Source |  |  |
|  | - Bidder should state if blades are reusable or disposable. |  |  |
|  | - Bidder should quote separately for the following blade sizes: - Sizes 3 (Curved), Qty. (1 each) - Size 4 (Curved), Qty. (1 each) - Special blade size dedicated for difficult intubation for adults must be quoted separately. Qty. (2 for the total quantity) - Other blade sizes and types to be priced separately |  |  |
|  | - Can be connected easily to the monitor through a connecting table with length of at least1.5 meters |  |  |
| 3 | Flexible intubation video endoscope (Adult) Qty. (1 each) |  |  |
|  | - With built in video camera (chip on the tip): CMOS technology that provides superior video image quality |  |  |
|  | - LED light source: built in |  |  |
|  | - Outer diameter of the distal tip: approximately 5.5 mm |  |  |
|  | - Working length ≥ 60 cm |  |  |
|  | - Deflection up/down |  |  |
|  | - Reusable & sterilizable type |  |  |
| 4 | Flexible intubation video endoscope (Pediatric) Qty. (1 for the total quantity) |  |  |
|  | - With built in video camera (chip on the tip): CMOS technology that provides superior video image quality |  |  |
|  | - LED light source: built in |  |  |
|  | - Outer diameter of the distal tip: approximately 2.9 mm |  |  |
|  | - Working length ≥ 50 cm |  |  |
|  | - Deflection up/down |  |  |
|  | - Reusable & sterilizable type |  |  |
| 5 | Protective case should be included |  |  |
| 6 | All standard accessories should be included |  |  |
| 7 | Option: Mobile trolley with basket for blades should be priced separately |  |  |
| 8 | Option: Video stylet to be priced separately if available |  |  |
| 9 | All consumables should be priced separately, prices should be fixed according to special terms |  |  |
| 10 | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |

| Item 21 | Adult / Paediatric ICU Ventilator | Qty. (8) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well known manufacturer |  |  |
| 2 | General Features |  |  |
|  | - Minimum 15 inch swivable TFT colour touch-screen mounted above the trolley |  |  |
|  | - Application: ICU, Adult & paediatric, Invasive & non-invasive Neonate option should be available and quoted separately (if not included) |  |  |
|  | - Should operate on central gas sources (oxygen and medical air) |  |  |
|  | - Ergonomically designed |  |  |
|  | - Construction: robust, corrosion resistant and spill proof |  |  |
|  | - Flow sensor: Bidder should state clearly if the flow sensor is disposable or reusable; price of the disposable flow sensor should be mentioned in the technical offer and fixed as per special terms |  |  |
|  | - Oxygen sensor: - Type of oxygen sensor should be stated clearly - Price & lifetime of oxygen cell should mentioned clearly in the technical offer and fixed as per special terms |  |  |
|  | - Open system for patient circuits |  |  |
|  | - Self-test capability |  |  |
|  | - Trolley: heavy duty stable design with medical grade castors and articulated flexible arm to hold patient circuit |  |  |
|  | - Gas hoses (Air and Oxygen): - Original make - Colour coded - Fitted with British Standard probes |  |  |
|  | - Humidifier: - Dual servo controlled type - Heated with automatic water filling - Temperature and fluid level alarms - All required accessories needed to operate the unit should be included |  |  |
|  | - Medical grade compressor (option to be priced separately): - Should operate automatically when the pressure of the central air line is decreased below a specific level - Should be mounted on the same trolley of the ventilator |  |  |
|  | - Upgradability: - Software upgradeable - Hardware upgradable |  |  |
| 3 | Operating Modes |  |  |
|  | - Assist / Control mode: volume and pressure |  |  |
|  | - SIMV modes: volume and pressure |  |  |
|  | - Spontaneous / CPAP mode |  |  |
|  | - Non Invasive Ventilation: Nasal CPAP or Mask |  |  |
|  | - Apnea back-up ventilation |  |  |
|  | - Automatic leak compensation |  |  |
|  | - Bi-level and high flow ventilation modes should be quoted separately (if not included) |  |  |
|  | - Adaptive ventilation mode: please specify function |  |  |
|  | - Additional modes and tools: any additional/ optional modes should be mentioned and quoted |  |  |
| 4 | Controls /Settings Ranges |  |  |
|  | - Tidal volume (ml) range to cover a minimum of 25-2000 |  |  |
|  | - Inspiration time, sec: to cover the range 0.2 -5 or better |  |  |
|  | - Peak flow (L/m): Minimum range 3-100 or better |  |  |
|  | - Inspiration pressure (cmH2O): Minimum range 5-80 or better |  |  |
|  | - Inspiration hold/plateau ≥ 2 sec |  |  |
|  | - Respiration rate (breaths/minute): Minimum range 5-80 or better |  |  |
|  | - I:E ratio: 1:8 to 4:1 or better |  |  |
|  | - FIO2: 21-100% |  |  |
|  | - PEEP (cmH2O):minimum range 1-30 or better |  |  |
|  | - Pressure support(cmH2O): Minimum range of 0-50 |  |  |
|  | - Nebulizer: built-in |  |  |
|  | - Trigger mechanism to be pressure or flow |  |  |
| 5 | Monitored / Displayed Parameters |  |  |
|  | - PEEP pressure |  |  |
|  | - Pressure data: Peak / plateau /mean |  |  |
|  | - Respiratory rate |  |  |
|  | - FIO2 % |  |  |
|  | - Expiratory minute volume |  |  |
|  | - Tidal volume |  |  |
|  | - Leak |  |  |
|  | - Flow |  |  |
| 6 | Patient Alarms Audible and visual alarms for monitored parameters (user adjustable) : |  |  |
|  | - Low / high FiO2 |  |  |
|  | - High/Low minute volume |  |  |
|  | - High/Low tidal volume |  |  |
|  | - High airway pressure |  |  |
|  | - High respiratory rate |  |  |
|  | - Apnea time |  |  |
| 7 | Equipment Alarms |  |  |
|  | - Gas supply loss |  |  |
|  | - Mains Power failure |  |  |
|  | - Ventilator inoperative |  |  |
|  | - Patient circuit disconnection |  |  |
|  | - Low battery |  |  |
|  | - Continuous audible alarm for at least one minute in case of mains power failure even when the battery is depleted |  |  |
| 8 | Electric Requirements |  |  |
|  | - Mains: 220 Volts A.C, frequency 50 Hz |  |  |
|  | - Backup Battery: - Back-up time for settings and operation: minimum of 30 minutes - Batteries should be easy replaceable - Battery type and specifications should be mentioned clearly, bidder should state clearly if the battery is available in the local market; if else it price of the battery should be mentioned in the technical offer and fixed as per special terms - Battery should be replaced in free of charge basis during the warranty period - Any required battery conditioner should be included |  |  |
| 9 | Preventive Maintenance - should be the sole responsibility of the supplier during the warranty period including all required service parts/ kits - It should be conducted as recommended by the manufacturer |  |  |
| 10 | Standard Accessories List all standard accessories required for commissioning the system and indicate if they are included in the main unit price |  |  |
| 11 | For the final list of offers having equal chances of winning the award, the awarding process will be based on the accumulative value of both the offered item and its’ running cost (flow sensors, oxygen cells, etc….) over a period of seven years as per special terms (Total cost of owner ship) |  |  |
| 12 | Optional Accessories List all optional accessories and price them separately including medical grade air compressor |  |  |
| 13 | Abroad service training for 1 biomedical engineer/ technician as per special terms |  |  |

| Item 22 | Patient Radiant Warmer, Adult & Paediatric | Qty. (8) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Application: for adult & paediatric patients in recovery, intensive care departments and hospital wards |  |  |
| 3 | Safe warming operation without damaging patient skin layers and eyes. |  |  |
| 4 | Fast heat operation with durable long life heating elements and protective cover to prevent hand burns |  |  |
| 5 | Mounted on antistatic height adjustable pole; pole to be made of heavy duty stainless steel or equivalent material with high quality castors and braking mechanism |  |  |
| 6 | Timer for setting of heating duration. |  |  |
| 7 | Adjustable heat intensity |  |  |

| Item 23 | Laryngoscope Set, 4 Blades, Adult & Pediatric | Qty. (11) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Light weight. |  |  |
| 3 | LED type. |  |  |
| 4 | Fibre optic plate illumination type. |  |  |
| 5 | To be made from heavy-duty stainless steel (Handle & Blades) |  |  |
| 6 | Standard size, non-rechargeable battery operation. |  |  |
| 7 | Each unit shall include: - Size (3 & 4) stainless steel curved blades - Size (1 & 2) stainless steel straight blades - Size (5) stainless steel curved blade to be priced separately |  |  |
| 8 | Durable case with sufficient padding. |  |  |

| Item 24 | ICU Patient Monitor | Monitors: Qty. (60) Central station : Qty. (6) Portable Monitor: Qty. (7) CO2 Module: Qty. (10) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements |  | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- | --- |
| A | Patient Monitor: From a reputable well-known manufacturer, durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |  |
|  | - ICU/CCU patient monitor |  |  |  |
|  | - Adult, paediatric and neonatal patient monitor. |  |  |  |
|  | - High resolution TFT-LCD touch-screen display of at least (1024 x 768) resolution and size not less than 15" |  |  |  |
|  | - Ability to display 6 waveforms simultaneously |  |  |  |
|  | - User friendly. |  |  |  |
|  | - Heavy duty with robust design. |  |  |  |
|  | - Configuration: modular type with a minimum of three slots for three plug-in modules that can be fitted simultaneously on a module rack/ slots. Or One slot/ docking station for one removable module that is capable to measure all parameters mentioned in (8 & 9 & 10). |  |  |  |
|  | - Multi parameter module to measure the following parameters: ECG, Respiration rate, NIBP, Body temperature, SPO2 (Masimo technology or any equivalent technology that rejects artifacts from low body perfusion and motion artifact) & IBP |  |  |  |
|  | - Additional IBP (outlet or separate module) to be quoted separately |  |  |  |
|  | - End tidal CO2, priced separately Qty.(10) |  |  |  |
|  | - Ability to measure 12 ECG leads, software should be included |  |  |  |
|  | - ST Histogram |  |  |  |
|  | - QT & QTc analysis |  |  |  |
|  | - Wall- mounting accessories supplied and approved by the manufacturer of the monitor. |  |  |  |
|  | - Leads off alarm |  |  |  |
|  | - Escalating level audio visual alarm that cannot be permanently silenced |  |  |  |
|  | - Original accessories (priced separately): |  |  |  |
|  | Accessory | Total required Qty. |  |  |
|  | - ECG cable (3 leads), with ECG trunk cable | 45 |  |  |
|  | - ECG cable (5 leads), with ECG trunk cable | 45 |  |  |
|  | - Reusable SpO2 sensors (adult), 3 meters long, Finger clip type | 90 |  |  |
|  | - Reusable SPO2 sensors (Paediatric), 3 meters long, Finger clip type | 30 |  |  |
|  | - Reusable multisite SPO2 sensors, 3 meters long | 10 |  |  |
|  | - Large adult cuff, reusable | 30 |  |  |
|  | - Regular size adult cuffs, reusable | 90 |  |  |
|  | - Small adult size cuff, reusable | 30 |  |  |
|  | - Paediatric cuff, reusable | 30 |  |  |
|  | - Neonate cuff, disposable | 100 |  |  |
|  | - Temperature probe, skin type, adult & Paediatric | 60 |  |  |
|  | - CO2 kit | 30 |  |  |
|  | - Accessories basket (supplied and approved by the manufacturer) | 60 |  |  |
|  | - All data shall be in real time (waveform and numeric). |  |  |  |
|  | - Software shall be of the latest version, and any service / settings password must be disclosed and perpetual |  |  |  |
|  | - Vendors to list all prices of optional software and hardware for offered system. |  |  |  |
|  | - Any software licenses must be listed and priced. |  |  |  |
|  | - Trend storage capbility to store patient data for at least 72 hours |  |  |  |
|  | - All available options and accessories must be offered and priced separately |  |  |  |
|  | - Prices of all accessories (including ECG cables, SPO2 sensors, cuffs, pressure transducers, temperature sensors etc.) should be priced separately and fixed as per special terms. |  |  |  |
|  | - Small modular portable monitors (Qty. 7) should be quoted; with built in rechargeable battery, multi-parameter module. Module should be compatible with the bed side monitor and should have the capability to store all critical patient information during transportation |  |  |  |
| B | Central Station Qty.(6) |  |  |  |
|  | - Bed licenses: - Qty. (2) Central stations: 10 bed licenses (Day Case) - Qty. (4) Central station: 8 bed licenses (Recovery) |  |  |  |
|  | - The central station is to be located at the nurse station. Bidders shall quote for a comprehensive offer including network and any required cabling, switches, etc….. The local area network (LAN) must be designed so that the failure of any bedside monitor does not affect the function or balance of the system. |  |  |  |
|  | - Central station must be supplied by the same manufacturer of the patient monitor, locally supplied central stations will be considered as non-conforming. |  |  |  |
|  | - The central station monitor shall display waveforms for the following parameters as obtained from each individual bedside monitor: ECG, NIBP, IBP, Body temperature, SPO2, CO2 |  |  |  |
|  | - Ability to display waveforms, numeric data, graphic displays, tabular displays, and calculations. |  |  |  |
|  | - Remote record initiation control for all waveforms. |  |  |  |
|  | - Audio / visual alarms with standard alarm / acknowledge / reset control. |  |  |  |
|  | - Digital display of alarm limits and status. |  |  |  |
|  | - Memory for alarm events to provide a minimum of 10 seconds before the event, and 15 seconds after the event. |  |  |  |
|  | - Alphanumeric data display for each parameter. |  |  |  |
|  | - High resolution TFT-LCD touch screen colour monitor minimum 19 inches |  |  |  |
|  | - Multi-channel strip-chart recorder or laser printer with the following features: |  |  |  |
|  | - Annotation of time, date, bed number, alarm status, and vital signs. |  |  |  |
|  | - Paper width minimum 50 mm. |  |  |  |
|  | - Printer to allow for manual initiation from either the central station or any bedside monitor. |  |  |  |
|  | - Clinical review software with the following features: |  |  |  |
|  | - Ability to store patient data for at least 96 hours. |  |  |  |
|  | - Detailed analysis of waveforms. |  |  |  |
|  | - Clinical review software capable of generating full documented report. |  |  |  |
|  | - HL7 licence |  |  |  |
|  | - All available options and accessories must be offered and priced separately |  |  |  |
| C | Patient Monitor Analyser: Offer must include one patient monitor analyser (priced separately) - From well-known manufacturers - One compact unit with the ability to measure the following: - ECG - Respiration Rate - Temperature - Non Invasive Blood Pressure (NIBP) - Invasive Blood Pressure (IBP) - Oxygen Saturation (SPO2) - Analyser should be compatible with patient monitors from at least the following manufacturers: Philips, GE & Nihon Kohden. - All accessories needed to run the analyser should be included |  |  |  |
| D | Abroad service training for at least 4 biomedical engineers/ technicians as per special terms |  |  |  |

| Item 25 | Spot-Check Vital Signs Monitor | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements |  | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- | --- |
| 1 | From a reputable well-known manufacturer, durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |  |
| 2 | The vital signs monitor should be used to rapidly & accurately spot check monitoring of the following vital signs parameters: |  |  |  |
|  | - Noninvasive blood pressure: with accuracy ±5mmHg or better |  |  |  |
|  | - Temperature measurement by one of the following technologies: - Welchallyn - Covidien |  |  |  |
|  | - SpO2 pulse Oximeter measurement with accuracy ± 2% or better |  |  |  |
|  | - Heart rate |  |  |  |
| 3 | LCD display of size not less 5.5 inches |  |  |  |
| 4 | To work on either direct ac power supply or through ac/dc adapter |  |  |  |
| 5 | Built in rechargeable battery: - Minimum of 3 hours operation when fully charged - Battery type and specifications should be mentioned clearly, bidder should state clearly if the battery is available in the local market; if else it price of the battery should be mentioned in the technical offer and fixed as per special terms |  |  |  |
| 6 | HL7 License or DICOM (Compatible with Hakeem Hospital Information System without any extra cost) |  |  |  |
| 7 | Original heavy duty mounting trolley on five castors and accessories basket |  |  |  |
| 8 | Original accessories (priced separately): |  |  |  |
|  | Accessory | Total required Qty. |  |  |
|  | - Adult cuff | 8 |  |  |
|  | - Pediatric cuff | 4 |  |  |
|  | - Large adult cuff | 4 |  |  |
|  | - SpO2 probe, reusable, adult | 8 |  |  |
|  | - SpO2 probe, reusable, pediatric | 4 |  |  |
|  | - Temperature sensor oral | 4 |  |  |
|  | - Temperature sensor cover | 400 |  |  |
| 9 | Price of accessories and consumables including temperature sensor covers must be fixed as per special terms |  |  |  |
| 10 | For the final list of offers having equal chances of winning the award, the awarding process will be based on the accumulative value of both the offered item and its’ running cost (temperature sensor cover and any other required consumable) over a period of seven years as per special terms (Total cost of owner ship) |  |  |  |
| 11 | Printer option to be quoted separately (if available) |  |  |  |

| Item 26 | Oxygen Flow Meter Wall Type /Single | Qty. (77) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction and reliable |  |  |
| 2 | Oxygen flow meter to provide accurate gas flow measured and controlled within a range of 1-15 L/min, for use in a variety of respiratory therapy clinical devices. |  |  |
| 3 | The flow meter gauge should be float in tube |  |  |
| 4 | Clearly labelled and colour-coded flow meter |  |  |
| 5 | Offer shall include 200-300 ml reusable humidifier mounted directly on the flow meter (priced separately) |  |  |
| 6 | Quick connection type: BOC. |  |  |
| 7 | - Accuracy: +/- 10% or better |  |  |
| 8 | Wall mounted. |  |  |

| Item 27 | Regulator Suction with canister, wall vacuum outlet | Qty. (77) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | A continuous vacuum regulator for use with all common suction therapy procedures to control suction from wall vacuum outlets. Durable, heavy duty construction and reliable |  |  |
| 2 | Compact, lightweight regulator to be made of strong, break-resistant material |  |  |
| 3 | Large suction pressure control knob |  |  |
| 4 | - Flow rate: 0 – 40 L/min. |  |  |
| 5 | - Gauge: 0-700 mmHg approximately |  |  |
| 6 | Suction set shall consist of: |  |  |
|  | - Quick connection type: BOC |  |  |
|  | - All required tubing |  |  |
|  | - Biological filter (should be priced separately, price should be fixed as per special terms) |  |  |
|  | - One autoclavable jar min of two litres capacity with wall bracket |  |  |
|  | - Overflow protection device with wall bracket if needed. |  |  |
|  | - Disposable suction bags should be quoted separately |  |  |

| Item 28 | Oxygen Regulator for Oxygen Cylinders/ Low Flow | Qty. (12) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction and reliable |  |  |
| 2 | To be used with oxygen cylinders. |  |  |
| 3 | Heavy duty metallic construction (stainless steel or brass). |  |  |
| 4 | At least one gauge to indicate pressure inside cylinder |  |  |
| 5 | Flowmeter: 1-15 LPM with reusable humidifier |  |  |
| 6 | Bull nose connection. |  |  |

| Item 29 | Pulse Oximeter, Fingertip | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, compact, heavy duty construction, reliable, ergonomic design and easy to use |  |  |
| 2 | Provides fast and accurate measurements of oxygen saturation and pulse rate in paediatric to adult patients |  |  |
| 3 | Battery operated model on regular size batteries that are available in the local market |  |  |
| 4 | Digital display of SPO2 and heart rate |  |  |
| 5 | SPO2 range 1-99 %. Or better |  |  |
| 6 | Pulse rate measurement range: 30 – 240 beats per minute (bpm) or better |  |  |
| 7 | Measurement accuracy not less than 97%. |  |  |
| 8 | Original carrying case to be included |  |  |

| Item 30 | Digital Scale, Adult & Paediatric | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Suitable for adult and paediatric patients |  |  |
| 2 | # It shall have digital readout of patient weight in kg |  |  |
| 3 | ## It should weight accurately to the nearest 100g or better |  |  |
| 4 | ## It should be made from sturdy plastic/steel material. |  |  |
| 5 | ## Maximum capacity should be 200 Kg / 440 lbs or better. |  |  |
| 6 | ## It should power up from AC voltage |  |  |

| Item 31 | Oxygen Cylinder, 2 Cubic Meters | Qty. (12) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Nominal contents 2 cubic meters approximately |  |  |
| 2 | Nominal cylinder pressure: 137 bar |  |  |
| 3 | Valve outlet connection: Bullnose according to BS standard |  |  |
| 4 | Material of construction: steel |  |  |
| 5 | Colour: according to BOC colour code system |  |  |
| 6 | Hand-wheel & safety valve |  |  |

| Item 32 | ECG Recorder System | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements |  | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- | --- |
| 1 | Durable, compact, heavy duty construction & reliable from well-known company |  |  |  |
| 2 | Microprocessor controlled with self-test and error code/message capability. |  |  |  |
| 3 | 12 lead ECG acquisition |  |  |  |
| 4 | Display: |  |  |  |
|  | - Size: At least 5.7 inch LCD screen |  |  |  |
|  | - Ability to display at least 6 ECG leads simultaneously and heart rate indicator |  |  |  |
| 5 | ECG interpretation capability with event marking function |  |  |  |
| 6 | Defibrillation protection: Type CF defibrillation proof |  |  |  |
| 7 | Filters: AC noise filters (50/60 Hz) and muscle artifact filters |  |  |  |
| 8 | Built in rechargeable battery: - Battery type and specifications should be mentioned clearly, bidder should state clearly if the battery is available in the local market; if else it price of the battery should be mentioned in the technical offer and fixed as per special terms - Battery should be replaced in free of charge basis during the warranty period |  |  |  |
| 9 | Digital Thermal Printer: |  |  |  |
|  | - High resolution print of ECG report including patient data and interpreted diagnostic aids. |  |  |  |
|  | - Recoding paper: Thermal paper of A4 size |  |  |  |
| 10 | Original heavy-duty mobile cart with articulated arm to hold ECG cable and at least one drawer |  |  |  |
| 11 | Original Accessories (priced separately): |  |  |  |
|  | Accessory | Total Required quantity |  |  |
|  | ECG patient cable (12 leads) | 6 |  |  |
|  | Full set of suction electrodes | 8 |  |  |
|  | Full set of limb clamps | 8 |  |  |
| 12 | Price of all accessories including ECG cables, suction electrodes, limb clamps and should be fixed as per special terms |  |  |  |
| 13 | HL7 License or DICOM (Compatible with Hakeem Hospital Information System without any extra cost) |  |  |  |
| 14 | Offer must include one ECG simulator (priced separately): - From well-known manufacturer - 12 Lead - Compatible with different ECG manufacturers - Complete with all standard accessories |  |  |  |

| Item 33 | Aneroid Sphygmomanometer/ Mobile Stand | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, compact, heavy duty construction & reliable, ergonomic design and easy to use |  |  |
| 2 | Large, easy-to-read scale of at least 12 cm in diameter and up to 300 mmHg range |  |  |
| 3 | Accuracy up to ± 3 mmHg |  |  |
| 4 | Latex free inflation system |  |  |
| 5 | Mobile stand: Sturdy, stable that is not easy to be tilted and easy to manoeuvre with at least 50 cm diameter base and 5 heavy duty castors |  |  |
| 6 | Large storage basket mounted on the trolley |  |  |
| 7 | The following original accessories must be included with each unit: |  |  |
|  | - Two durable regular size adult cuffs |  |  |
|  | - One durable paediatric cuffs |  |  |
|  | - Two inflation bulbs |  |  |
|  | - One spiral tubes of at least 1 meter long each |  |  |
| 8 | All parts (tubing, valve, bulb…etc.) must be original from the original manufacturer. |  |  |
| 9 | Kindly quote separately for all other sizes of cuffs including obese adult cuff. |  |  |

| Item 34 | Digital Thermometer | Qty. (12) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Fast: quickly captures accurate readings on all patient types |  |  |
| 2 | Convenient storage for 25 probe cover |  |  |
| 3 | Large LCD displays temperatures in Celsius |  |  |
| 4 | Oral probe type, rectal type to be priced seprately |  |  |
| 5 | Last temperature recall button |  |  |
| 6 | Interchangeable, removable probe covers to help reduce the risk of cross-contamination |  |  |
| 7 | Probe covers to be priced separately |  |  |
| 8 | For the final list of offers having equal chances of winning the award, the awarding process will be based on the accumulative value of both the offered item and its’ running cost (probe covers) over a period of seven years as per special terms (Total cost of owner ship) |  |  |
| 9 | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |

SPECIAL TERMS

- Offers not complying with any of the special terms or the technical specifications shall be considered non-conforming with tender requirements.

- Any vendor providing FORGED documents shall be disqualified from the current tender and banned from participating in any future RMS tenders.

- Wherever term “based” is mentioned it refers to the country where the manufacturing company is founded & established.

- All equipment must be the most recently released model/version which is equal to or higher than the range of the specifications of the required system (low, mid or high) and equal to or higher than the level of technology and required options mentioned in  the technical specifications.

- Required certificates (must be submitted with the technical offer):

- FDA clearance (510K) for equipment of USA based companies.

- MDR (EU) 2017/745 certificate for equipment of EU (European Union) based companies.

- UKCA certificate For Equipment of Great Britain based companies (England, Scotland, Northern Ireland, and Wales).

- ARTG (Australian Register of Therapeutic Goods) certification or approval for Australian and New Zealand based companies.

- Establishment License with the relevant Device License issued by the Therapeutic Products Directorate for Equipment of Canadian based companies.

- PMDA (Pharmaceuticals and Medical Devices Agency) certification or approval for equipment of Japanese based companies.

- Swissmedic (Swiss agency for therapeutic products) certification or approval for equipment of Swiss based companies.

- Norwegian Medicines Agency certification or approval for equipment of Norwegian based companies.

- Only for class Ι medical equipment manufactured by companies based in one of the countries mentioned above, submission of either one of the certificates mentioned above or a free sale certificate in any of these countries shall be accepted.

- For the following equipment:

- All X-ray equipment, MRI, ultrasound, and nuclear medicine systems (regardless where the manufacturing companies are based).

- Where the manufacturing companies are based in other origins than the mentioned in terms 2.1 - 2.8.

The following are required:

- At least two of the certificates mentioned above, one of which has to be FDA clearance (510K) (Only for class Ι medical equipment submission of certificates mentioned in 2.9 shall be accepted).

- Evaluation certificate from the Royal Medical Services for the same offered model with at least 80% passing grade.

If the evaluation is not applicable (based on purchasing committee perspective) bidder should submit a list of installation basis of the same offered model and/ or previous models in at least two of the following hospitals (King Hussein Cancer Center, National Center for Diabetes Endocrinology and Genetic Diseases, Jordan University Hospital or King Abdullah University Hospital) with at least three years of operation, list should include: Name of hospital, Model installed, Quantity, and date of installation.

The purchasing committee has the right to officially contact any of these hospitals and disqualify any offer where the feedback is negative in operation, after sales service or local agent performance.

- The vendor is responsible to ensure through official  documents that classified medical devices are manufactured in conformity with applicable quality system standards (ISO, IEC); (the international quality systems standard for medical device manufacturing ISO 13485) or ISO9001.

- With each offer, bidders must provide a formally endorsed document issued by the manufacturer stating that the bidder is the sole certified agent for the offered item.

- In all of the above cases (except 2.12) certificates must be formally endorsed by JFDA.

- Any vendor not submitting all required certificates will be eliminated.

- Official website of the manufacturer must be provided. The website must clearly demonstrate the history profile and manufacturing features of the company including the same offered model and its brochure.

In case of inconsistency between the information on the website with the provided information of the local agent regarding any offered product, the committee has the right to eliminate that product from the awarding process.

- Offered items should be from reputable well known manufacturers and excellent experience in the field and shall have multiple installations of the same offered model and/ or previous models in RMS main hospitals with at least two years of operation and excellent experience in operation, after sales service & local agent performance; otherwise the purchasing committee has the right to request any of the following:

- An evaluation certificate as mentioned in term 2.10.b.

- A sample of the same offered model at any time during the purchasing process, the sample unit shall be delivered within three weeks from the date of a written notification; any vendor rejects to deliver a sample unit will be eliminated. Evaluation duration will be determined by the purchasing committee.

Any offered item fail in the evaluation/assessment process will be rejected

- Vendors must specify the origin of the offered items and accessories in the technical offer.

- Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first half of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period shall be rejected.

- Goods are to be brand new; manufacturing date of the delivered products should not exceed 18 months from the date of the final award.

- Delivery period should be mentioned clearly; the purchasing committee has the right to reject any offer with delivery time exceeding 4 months.

- Warranty:

- Offers must include a full warranty for a period of a minimum of 24 months from the date of installation or 30 months from the date of receiving the items at the agreed location mentioned in the final order (whichever comes first).

Warranty must include corrective & preventive maintenance activities as per manufacturer recommendations including:

- Required spare parts (free of charge)

- Labour

- Hardware

- Software

- Rechargeable batteries

At the end of the warranty period, The supplier commits to implement final inspection of the submitted goods and submit the reports signed by the site chief engineer stating that the equipment are working properly as well as all preventive maintenance reports during the warranty period.

Warranty & installation are excluded for items mentioned in Attachment no.1 the purchasing committee has the right to exclude any other item not listed in the attachment and does not required installation and warranty.

- In the case where a delay in installation has occurred as a result of the supplier's dereliction and has exceeded a period of one month from the date of receiving the items, the approved warranty period will automatically be considered as 24 months from the date of installation.

- If at any time during the warranty period the item becomes inoperative due to a technical fault the item must then be repaired by the supplier /local agent within a period of 5 working days from written notification, warranty will be extended according to downtime period.

- If the delay exceeds 30 days the supplier must replace the item with a new identical functioning one (within the same delivery period mentioned in the final offer). In case the item was replaced by a new one, the warranty period mentioned in 9.a) above will start from the installation and commissioning date of the new item.

- Local agent/ supplier is committed to transport and install any awarded item during the warranty period on free of charge basis to any location inside the country whenever required by Royal Medical Services, this should not include pre-installation requirements in the new site.

- All offered items (main unit) should be fully designed, manufactured, and labelled by their real original manufacturer in which all related testing, research, development and approvals went through.

Any relabelled products for the main unit (white-label manufacturing, OEM, or repackaging) are rejected.

- One set of operation manual(s) and one set of service manual(s) including schematics and a spare-part list must be delivered with each unit, CD/DVD is acceptable. For large tenders, a certain agreed percentage of manuals per item may be agreed upon.

- Where applicable, pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include removal of old system(s), any civil work, electrical work or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new system(s).

- Power requirements: where applicable either single phase 220V, 50Hz or 3-phase 380V. Systems with external transformers are considered conforming only if clearly stated in the technical specifications.

- Technical offers must include clear original technical brochures/catalogues for all offered items.

- Offers must include fully detailed technical offers and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s), full technical description/specifications and any accessories or options included in the offer.

Offer must clearly indicate the origin of the offered item and the country where the manufacturer of the offered items is based.

- Compliance sheets must be as per the tabular format of the technical specifications in the tender documents, listing the required specifications on one column and a  Yes  or NO  response to each point in the adjacent column, with reference to page and line numbers in the relevant technical brochure. Offers not complying with this term shall be rejected.

- Qualifications and after sales service: The technical bid should contain all the necessary documents (training certificates) to prove the capability of the bidders for supplying and installing a trouble free equipment meeting the quality standards and technical specifications of the manufacturer and the ability of the bidders for providing efficient after sales service conducted by authorized certified biomedical service engineers with minimum 2 years experience in the same field.

- Accessories and consumables:

  - Any accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately.

All offered accessories and consumables must be approved by the manufacturer.

  - Technical offers must include a priced list for all accessories and consumables related to the required equipment as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) with prices fixed for a period of five years from the date of installation and commissioning with a maximum annual increase of 2%, any essential item not listed will be considered free of charge.

  - Accessories and consumables must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Where applicable, a start-up kit of accessories and consumable items must be provided with each system on a free-of-charge basis.

- Spare Parts:

  - Technical offers must include a comprehensive and priced list of all spare parts related to the awarded equipment (including rechargeable batteries) as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) valid for a minimum period of five years with a maximum annual increase of 2%, commencing at the end date of the warranty period, any essential item not listed will be considered free of charge.

  - Spare parts must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Prices of spare parts should be reasonable and will be taken into consideration during the purchasing process; the purchasing committee has the right to eliminate any offered item with unreasonable high prices of spare parts.

  - Delivery period of required spare parts should not exceed 2 months from the date of the final order.

- Spare parts, consumables and accessories availability must be guaranteed for a minimum period of ten years starting from the date of installation and commissioning.

- Local agent is committed to sign a service contract with Royal Medical Services to perform all relevant maintenance for the awarded equipment (whenever required by DRMS within a period that does not exceed one year from the end date of the warranty period and if mentioned clearly in the specifications) without exceeding the original mentioned cost and as per mentioned terms within minimum required period.

- Tender Awards:

  - For the final list of offers having a chance of winning the award, the awarding process shall be based on the accumulative value of both the offered item and its’ running cost (Total Cost of Ownership) over a period of seven years from the date of installation and commissioning. Only offers with the lowest total cost of ownership over a period of seven years from the date of installation and commissioning shall qualify for the award.

  - Running cost includes the value of consumables, accessories needed to operate the system over the same period as well as the cost of any service contract (where applicable).

- For PC/Laptop based systems:

  - Complete restoration medium (CD/DVD/etc.) of the operating system and the application software must be supplied.

  - Purchased licenses, software keys & dongles must be provided at any time during and after the warranty period on free of charge basis whenever required by DRMS for a minimum period of ten years starting from the date of installation and commissioning.

  - Where locally supplied computers, laptops & printers are offered, the offered model should be from well known manufacturer.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores, installation, pre-installation (if needed), training, commissioning, warranty and bringing the equipment into service.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- a. DRMS has the right to increase the awarded quantities by a percentage not exceeding 35% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

b. DRMS has the right to decrease the awarded quantities by a percentage not exceeding 50% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- a. Bidder is not allowed to submit more than one offer for the same item whether that’s solely or in coalition or partnership with other bidders.

b. Bidder is allowed to include within their offer optional alternatives for the same offered item from the same manufacturer.

- The supplier must furnish DRMS with a guarantee stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelve months from the date of final acceptance of the equipment by DRMS.

- Training:

  - For items where abroad service training courses are required in technical specifications, offers must include a certified service training program for at least 3 working days at a reputable center abroad recognized by the manufacturer for at least one biomedical engineer or biomedical technician; all costs inclusive, air tickets, , boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - For items where abroad user  training courses are required in technical specifications, offers must include a certified operator training program at a reputable center abroad recognized by the manufacturer for at least one operator; all costs inclusive, air tickets, boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - The period of the training courses must be according to the manufacturer’s program excluding traveling days and must be stated clearly in the technical offer.

  - Training Programs must conform to the following standards:

- User training must comprise understanding and use of operation manual(s), correct and safe operation of the equipment, as well as user preventive maintenance and calibration.

- Service training must comprise: theory, understanding and use of service manual(s), calibration, preventive maintenance procedure, and practical troubleshooting and repair exercises, and must be conducted by professional instructors employed or authorized by the system manufacturer.

- Service training must be conducted on a system of identical make, model, and configuration to that purchased by DRMS, and designated by the manufacturer or the local agent for training purposes.

- Certificates must be endorsed and officially sealed by the system manufacturer, legally empowering trainees to engage in user and service activities according to operation and service manual(s).

- Where applicable, offers must include an on-site user and service training.

30. For offers submitted in Jordanian dinars, Payment will be either by wire transfer or by cheque after final acceptance of goods. Any other way of payment will be rejected.

Attachment 1

| # | Equipment |
| --- | --- |
| 1 | Adjustable Auto Pipettes |
| 2 | Resuscitation Bag |
| 3 | Laryngoscope Set |
| 4 | Oxygen Flow meter wall type/ single |
| 5 | Regulator Suction with canister, wall vacuum outlet |
| 6 | Oxygen Regulator for Oxygen Cylinders |
| 7 | Pulse Oximeter, Finger type |
| 8 | Oxygen Cylinder |
| 9 | Doppler, portable |
| 10 | Diagnostic set, Portable |
| 11 | Direct Ophthalmoscope, Portable |
| 12 | Otoscope, Portable |
| 13 | Air Mattress System, homecare |
| 14 | Stethoscope |
| 15 | Aneroid Sphygmomanometer |
| 16 | Video Assisted Laryngoscope, portable |
| 17 | Scale, Manual |
| 18 | Wood's Light |
| 19 | Cough Pressure, Normal Saline |
| 20 | Rehabilitation Walking Parallel Bars, non-powered |
| 21 | Therapy Mat |
| 22 | Medical Ball All Size |
| 23 | Dumbbells Rack with complete set of dumbbells |
| 24 | Crutches |
| 25 | Shoulder wheel |
| 26 | Mobile Mirror |
| 27 | Cuff Weights |
| 28 | Walker, different sizes |
| 29 | Patient Elbow Stick |
| 30 | Rehabilitation Training Ladder |
| 31 | Rehabilitation Suspension Frame |
| 32 | Exercise Band All Size (Theraband) |
| 33 | Lens trial set |
| 34 | Wheel Chair |
| 35 | Commode Chair |
| 36 | Bassinet (Baby Cot) |
| 37 | Resuscitation Cart (Crash Cart) |
| 38 | Medication Cart |
| 39 | Cart, Drawers |
| 40 | Examination Couch, Manual |
| 41 | Gynaecology Examination Table, Manual |
| 42 | Examination Table, Neonates, Manual |
| 43 | Intravenous Pole, Mobile Stand |
| 44 | DDA Cabinet |
| 45 | Stainless steel Multipurpose Trolley |
| 46 | Cabinet, Instrument, Operation Theatres |
| 47 | Dressing Cart |
| 48 | Stainless steel wire shelving unit |
| 49 | Paper Trolley |
| 50 | Stainless Steel Sink (Clean up counter) |
| 51 | Scopes Cabinet |
| 52 | Mayo Table |
| 53 | Table, Instrument |
| 54 | Stool, Adjustable, Doctor |
| 55 | Stool, Adjustable, Operation Theatres |
| 56 | Carts, linen/laundry, soiled, Double |
| 57 | Closed distribution trolley |
| 58 | Step Ladder, Conductive, Double |
| 59 | Step, Surgeon, Single |
| 60 | Kick Bucket |
| 61 | Mobile Stand for Oxygen Cylinder |
| 62 | Stainless Steel Wire Basket, 1 STU |
| 63 | Cart, Plaster |

Annex 3:

A:  Neuro-Spine & Ortho-Spine

| Item A-1 | Neurology Surgical Endoscopy Tower | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| A | Durable, heavy duty and reliable neurology surgical tower from well-known manufacturer with at least the following specifications: |  |  |
| B | 4K Camera head: |  |  |
|  | - Compact, lightweight, easy to grip. |  |  |
|  | - Resolution: a minimum of 3840*2160. |  |  |
|  | - Gas sterilizable, Autoclavable is preferable. |  |  |
|  | - Focal distance and lens zooming. |  |  |
|  | - Immersible in disinfections liquid. |  |  |
|  | - Universal coupling mechanism to fit all standard endoscopes. |  |  |
|  | - Camera cable approx. 3m. |  |  |
| C | 4k Camera Controller (console): |  |  |
|  | - Clear function buttons and indicators. |  |  |
|  | - Hygienic. |  |  |
|  | - Digital display. |  |  |
| D | Light Source (either standalone device or integrated on camera control unit): |  |  |
|  | - LED Light source (light intensity should be equivalent to Xenon lamp 300W). |  |  |
|  | - Working hour ≥ 30,000 hours. |  |  |
|  | - Automatic and manual adjustment of light intensity. |  |  |
|  | - Front panel hygienic with convenient indicators for ease of operation. |  |  |
|  | - Flexible fibre optic cable preferably Autoclavable and not less than 230 mm. |  |  |
| E | LCD / LED medical grade Monitor: |  |  |
|  | - 4K Medical grade type that is supplied by the same manufacturer. |  |  |
|  | - Monitor size ≥ 27 inch. |  |  |
|  | - Resolution ≥ 3840*2160. |  |  |
|  | - To be compatible with the camera controller. |  |  |
|  | - DVI or HDMI output. |  |  |
| F | Archiving system: |  |  |
|  | - Recording images and videos. |  |  |
|  | - Has internal Hard Disk not less than 1 Terra |  |  |
|  | - Ability to record, archive and streaming |  |  |
|  | - All standard settings should be included. |  |  |
| G | Inflow/outflow pump: |  |  |
|  | - Touch screen ensure safe and precise adjustment of the set values |  |  |
|  | - parameters display, easy to read and arranged clearly parallel to one another allow the user to monitor the current actual and set values of all unit parameters at any time |  |  |
|  | - audible alarms in case of malfunction |  |  |
|  | - Irrigation Pressure: adjustable up to 150 mmHg or better |  |  |
|  | - Irrigation Flow Rate: adjustable up to 0.5 L/M or better |  |  |
| H | Telescopes & Instruments |  |  |
| H-1 | Interlaminar Endoscopic Surgical Set: Complete set (complete manufacturer catalogue list with sterilization container) to be used for interlaminar decompression or deherniation |  |  |
| H-2 | Transforaminal Endoscopic Surgical Set Complete set (complete manufacturer catalogue list with sterilization container ) to remove a herniated disc or to treat a foramen stenosis |  |  |
| I | Shaver drill system |  |  |
|  | - Dedicated for spine endoscopic surgeries |  |  |
|  | - High torque motor to remove soft and hard tissue and soft bones. |  |  |
|  | - Light weight and low vibration for increased surgeon tactile feel. |  |  |
|  | - Forward Speed of 6000 RPM or greater |  |  |
|  | - Modes: Forward / Reverse can be controlled in console and in foot switch. |  |  |
|  | - Continues variable speed foot switch. |  |  |
| J | Radiofrequency Ablation System |  |  |
|  | - Low temperature Bi-polar Radio frequency technology. |  |  |
|  | - Dedicated for spine surgeries to rapidly resects and sculpts soft tissue and coagulate bleeding vessels. |  |  |
|  | - Ablation and coagulation adjustment with different settings |  |  |
|  | - Foot switch |  |  |
| K | All accessories and consumables needed for operating all of the above components of the system (including radiofrequency ablation system) must be listed and priced separately for 200 cases, prices should be fixed as per special terms and will be taken into consideration during the purchasing pricing process (total cost of ownership) as per special terms. |  |  |
| L | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for the subcomponents for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |

| Item A-2 | Surgical Microscope, Neurosurgery | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # A floor mounted dedicated surgical microscope for use in neuro-spine surgeries. |  |  |
| 2 | # Magnification system that features motorized 1:6 zoom |  |  |
| 3 | # Binocular Tube : |  |  |
|  | - Three binocular tubes with apochromatic eyepieces for (Main, face to face assistance and lateral assistance) |  |  |
|  | - Tiltable tubes 0˚ - 180˚ |  |  |
|  | - Main eyepieces with at least 12.5 x magnification |  |  |
|  | - Assistant eyepieces with at least 10 x magnification |  |  |
| 4 | # Working distance with a range of 225- 450 mm or wider |  |  |
| 5 | # Auto Balance |  |  |
| 6 | # Manual backup of auto balancing of all axes |  |  |
| 7 | # Multifunctional 2 handles with programmable buttons |  |  |
| 8 | # X-Y movement ability: Microscope should include X-Y coupling for fine adjustment of microscope head |  |  |
| 9 | # Integrated Camera: HD resolution |  |  |
| 10 | # Diopter balance ( ± 5 dpt) |  |  |
| 11 | # 5-ALA Fluorescence (optional to be priced separately) |  |  |
| 12 | # Should be compatible and integrated with navigation system from well known manufacturers (Brainlab, Medtronic, Stryker) and the one purchased in item (A-3), bidder should mention the compatibility list, all hardware , software and licenses should be included |  |  |
| 13 | # Arm to reach not less than 1600 mm |  |  |
| 14 | # At least 21" Medical grade touch LCD monitor |  |  |
| 15 | # Electromagnetic braking system |  |  |
| 16 | # Two individual and adjustable lamps with minimum 300W Xenon. (Individual power supply for each lamp is preferable) |  |  |
| 17 | # Automatic lamp change when bulb is defective or failed |  |  |
| 18 | # Original digital video system with archiving capability on DVD or USB memory that is supplied and approved by the same manufacturer of the main system |  |  |
| 19 | # Should be compatible with HAKEEM hospital information system, all hardware, software and licenses should be included |  |  |
| 20 | # Foot brake for locking of moveable castors |  |  |
| 21 | # Switches and controls should be sealed against penetration of liquids |  |  |
| 22 | # Any other additional features available in the system are to be mentioned for comparison purposes |  |  |
| 23 | # All accessories and options must be listed and priced separately. |  |  |
| 24 | # A list of standard accessories for the offered model |  |  |
| 25 | # Abroad service training for at least 1 biomedical engineer/ technician as per special terms |  |  |

| Item A-3 | Intra-Operative Navigation System, Spine | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Intra-Operative Navigation System is required from a reputable well known company that has an excellent experience in this field and has multiple installation bases in many countries in Europe/ USA in addition to Jordan, a list of installation bases should be included in the technical offer |  |  |
| 2 | Application: Spine procedures for open and minimal invasive surgeries |  |  |
| 3 | System components: - Separate or standalone viewing & tracking carts. - Computer: - High speed CPU - At least 1 TB SSD - At least 32 GB RAM - All needed accessories should be included - DICOM compatible to transfer MRI and CT images to navigation system from the available RIS network, all licenses should be included - Ability to download MRI/ CT images from USB and CD/ DVD - High resolution (Full HD) medical grade monitor ≥ 27 inches |  |  |
| 4 | Tracking System: - Infrared camera for optical wireless marker tracking (or equivalent), suitable for spinal procedures for both adult and paediatrics - Bidders are required to specify and quote separately all needed disposables |  |  |
| 5 | Spinal Image registration: - Accurate, simple and fast - Anatomical landmark or trace registration for open spinal procedures - Automatic spinal registration: must be compatible and integrated with 3D C-Arm manufactured by at least the following (Siemens, GE, Ziehm) and purchased in item (A-5), all software, hardware and licenses should be included - Ability to re-register the patient intra-operatively in-case of patient or reference frame movement |  |  |
| 6 | Surgical Planning Software (Spinal): |  |  |
|  | - User friendly operating system |  |  |
|  | - Automatic image fusion software: - Fast and exact image fusion - Supports data of CT, MR (T1, T2 and MRA), PET CT and SPECT |  |  |
|  | - Segmentation of tumor and 3 D modelling |  |  |
|  | - Layout axial, sagittal, coronal, trajectory view, probe eye view & 3D view |  |  |
|  | - Virtual tool-tip extension |  |  |
|  | - Multiple trajectory planning for biopsy procedures |  |  |
|  | - Zoom function and adjustable grey level windowing |  |  |
|  | - Storage of screen shots for scientific documentation |  |  |
|  | - Spine screws planning length and diameter |  |  |
|  | - Any additional planning software features to be mentioned and priced separately |  |  |
|  | - Pre-planning of multiple trajectories (target/ entry point) |  |  |
|  | - Overview of all patient images for quick image selection |  |  |
|  | - Fast visualization of volumetric information in axial, coronal and sagittal planes |  |  |
| 7 | Universal instrument tracking adaptors, hardware & software (quoted separately) for the identification, visualization and calibration all standard instruments in the theatre (open system) of any additional surgical tool: - Spine screwdrivers: open to calibrate any spine system, spine tap - Spine AWL - Jamshidi needle |  |  |
| 8 | Biopsy Instrument Sets (priced separately) |  |  |
|  | - Biopsy aiming device kit |  |  |
|  | - Biopsy probe |  |  |
| 9 | Spinal Instrument Sets (priced separately) |  |  |
|  | - Open & MIS patient reference unit kit |  |  |
|  | - Sterile pointer |  |  |
|  | - Navigated AWL |  |  |
|  | - Navigated lumbar and thoracic probe (great shaft) |  |  |
|  | - Navigated cervical probe |  |  |
| 10 | Microscope interface system (quoted separately with all needed hardware and software) Compatible with at least Zeiss, Leica and the purchased model in item (A-2) |  |  |
| 11 | Disposables: all needed disposables for the following procedures should be priced separately and fixed as per special terms and will be taken into consideration during financial comparison (total cost of ownership) as per special terms : - 30 Spinal procedures - 30 Shunt procedures - 50 Biopsy procedures |  |  |
| 12 | Qualifications and after sales service: The technical bid should contain all the necessary documents (training certificates) to prove the capability of the bidders for supplying and installing a trouble free equipment meeting the quality standards and technical specifications of the manufacturer and the ability of the bidders for providing efficient after sales service conducted by at least one certified biomedical engineer/ technician |  |  |
| 13 | All other options should be listed and priced separately including cranial software with all required hardware requirements. Prices should be fixed for future upgradability as per special terms |  |  |

| Item A-4 | Intra Operative Neuromonitoring System for Spine Surgery | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Amplifier Box |  |  |
|  | - Number of Channels: 16 or 32Channels |  |  |
|  | - Band Width: ≈ 0.4Hz - 7KHz |  |  |
|  | - Sample Frequency: ≈ 25KHz |  |  |
|  | - Sensitivity: ≈ 0.01μV/Div to 30mV/Div |  |  |
|  | - Filters: High Pass: 0.1Hz-500Hz Low Pass: 10Hz-7000Hz Notch: 50 Hz, 60Hz or None |  |  |
|  | - Common Mode Rejection Ratio (CMRR): ≥ 120dB |  |  |
|  | - Noise Voltage: ≤ 2.0μVP-P |  |  |
|  | - A/D Converter: ≥ 24bit |  |  |
|  | - Input Impedance: ≥ 3000MΩ |  |  |
|  | - Scan Speed: ≈ 0.1-30000 mS/D, ≥20 grade adjustable |  |  |
|  | - Impedance Test: Real-time impedance of recording electrode scan be detected |  |  |
|  | - Low level stimulator: ≈ 0-8 mA, step value: 0.01 mA |  |  |
| 2 | Audio and Visual Stimulators |  |  |
|  | - Audio: Stimulate type: short sound/ pure sound/ white noise, Output: L/R/ L+R Max short sound intensity ≤ 125dB Max pure sound intensity ≤ 125dB Max white noise intensity ≤ 125dB. Stimulate Frequency: 0.1 Hz - 100 Hz or better |  |  |
|  | b) Goggles: Pulse Rate: ≈ 1-30Hz, Output: L/R/L+R Stimulate intensity: 1-15 or better |  |  |
| 3 | Software Modules: |  |  |
|  | - Free EMG |  |  |
|  | - Trigger EMG |  |  |
|  | - Upper SEP |  |  |
|  | - Lower SEP |  |  |
|  | - Trans cranial MEP |  |  |
|  | - TOF: Train of Four |  |  |
|  | - Pedicle Screw |  |  |
|  | - VEP: Visual Evoked Potential |  |  |
|  | - BAEP: Brain Auditory Evoked Potential |  |  |
|  | - Motor Mapping |  |  |
|  | - Speech Mapping |  |  |
|  | - D-Wave |  |  |
| 4 | Accessories and consumables: The system should accept the following from different manufacturers (open system), bidders are required to quote separately: |  |  |
|  | - Subdermal needles Qty. (100) |  |  |
|  | - Ground needles Qty. (50) |  |  |
|  | - Cork screws Qty. (2) |  |  |
|  | - Monopolar stimulation probe, reusable Qty. (5) |  |  |
|  | - Bipolar stimulation probe, reusable Qty. (5) |  |  |
|  | - E.M.G endotracheal tubes (optional to be priced separately) |  |  |

| Item A-5 | 3D C-Arm Imaging System | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | A state of the art digital Mobile three dimensional C-Arm fluoroscopy system is required to provide excellent image quality and minimal dose in neurosurgeries. The system must comprise sufficiently sophisticated hardware and software platforms to be able to host a wide range of real-time and post processing image enhancement techniques, the system must meet or transcend the following specifications: |  |  |
| 1 | Motorized lateral and orbital rotation in 3 dimensions (Coaxial, Coronal & Sagittal) for spinal and neurosurgical interventions |  |  |
| 2 | Must be compatible and integrated with navigation systems manufactured by at least (Brainlab, Medtronic, Stryker, Navient) and the one purchased in item (A-3), all software, hardware & licenses should be included |  |  |
| 3 | ## Physical Configuration: Two wheeled units with smooth manual steering system: one supporting the C-Arm and console and the other supporting monitors, image processing, recording devices etc. |  |  |
| 4 | ## System backbone: Microprocessor control over all subsystems with an elaborate self-test and error code scheme. |  |  |
| 5 | Control Console: hygienic touch type controls with easy to read, user friendly annotation. |  |  |
| 6 | X-Ray Tube: |  |  |
|  | - Rotating anode type. |  |  |
|  | - Anode heat Capacity, kHU ≥ 300 KHU |  |  |
|  | - Anode cooling, kHU/min ≈ 50 kHU/min |  |  |
|  | - Dual focus, size: ≤ 0.6 each |  |  |
| 7 | X-Ray Generator: |  |  |
|  | - Ripple-free high-frequency converter type. |  |  |
|  | - Power rating, kW@100 kVp: ≥ 15 kW. |  |  |
|  | - Radiographic Mode: - kV range: 40 to 120 kV or better in steps of 1 kV. - mAs range: up to 60 mA or better - Automatic exposure control |  |  |
|  | - Fluoroscopic Mode: - kV range: 40 to 110 kV or better - mA range: Up to 8 mA or better - Pulsed fluoroscopy: 1 to 25 pulses per second or better - Digital snapshot mode |  |  |
| 8 | Flat panel detector |  |  |
|  | - Size not less than (30 cm × 30 cm). |  |  |
|  | - Pixel size ≤ 200 µm or less |  |  |
|  | - CMOS flat panel detector type |  |  |
|  | - DQE ≥ 70% |  |  |
| 9 | TV Monitor: |  |  |
|  | - Dual high resolution displays of at least 18 inches coloured LCD monitors, or one 32 inches coloured LCD monitor |  |  |
|  | - Mounted on an articulating arm. |  |  |
| 10 | Image Acquisition: |  |  |
|  | - Image matrix: ~ 1024x1024. |  |  |
|  | - Collimation: Dual Leaf or Iris type. |  |  |
|  | - Image reversal and rotation capability. |  |  |
|  | - Magnification (detector zoom) |  |  |
| 11 | Image Processing and Storage: |  |  |
|  | - Image storage matrix: ~ 1024x1024. |  |  |
|  | - Patient data registration system. |  |  |
|  | - Last image hold. |  |  |
|  | - Cine replay. |  |  |
|  | - Digital subtraction. |  |  |
|  | - CD/DVD R archiving. |  |  |
|  | - Full DICOM 3.0 connectivity. |  |  |
| 12 | C-Arm: |  |  |
|  | - Free Space: 70cm |  |  |
|  | - Depth: 60cm |  |  |
|  | - Vertical Travel: 40cm |  |  |
|  | - Horizontal Travel: 20cm |  |  |
|  | - Orbital rotation ≥ 120˚ |  |  |
|  | - Pivot rotation ≥ ± 200˚ |  |  |
|  | - Panning motion ≥ 10˚ |  |  |
| 13 | A list of standard accessories for the offered model (optional). |  |  |
| 14 | Offer must include the price for a comprehensive service contract (including all replacement parts and x-ray tubes) valid for a period of (8) years starting at the end of the 24 month warranty period |  |  |

| Item A-6 | Neurosurgical Drill System | Qty. (1 set) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # High speed motor assembly with cable (RPM ≈ 75,000 or higher) Qty. (2) |  |  |
| 2 | # Console with LCD screen to control speed & direction |  |  |
| 3 | # Hand piece attachments: Straight, angled, long & short should be quoted separately Qty. (10) |  |  |
| 4 | # Accra Cut, Reusable, different sizes should be quoted separately Qty. (6) |  |  |
| 5 | # Footswitch |  |  |
| 6 | # Sterilization tray, with cover, for Hand piece and drill accessories Qty. (2) |  |  |
| 7 | # Drill Bits, all shapes and types should be quoted separately and fixed as per special terms and will be taken into consideration during financial comparison (total cost of ownership) as per special terms # Qty. (350) |  |  |

| Item A-7 | Sequential Air Compression System | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Intermittent pneumatic compression system to prevent deep vein thrombosis (DVT) |  |  |
| 2 | Latest model |  |  |
| 3 | LCD coloured touch display with at least 3 inch size |  |  |
| 4 | Light weight and small in size with carrying handle |  |  |
| 5 | Liquid proof design with at least IPX2 rating |  |  |
| 6 | Low noise operation |  |  |
| 7 | All standard accessories must be included (tubing and sleeves) and must be priced separately with all available sizes and types (leg and foot) |  |  |

| Item A-8 | Bone Mill | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Electrical type |  |  |
| 2 | Complete with all required control units and cables |  |  |
| 3 | Bench-top Small size unit |  |  |
| 4 | Produces appropriately sized bone particles without exposure to heat |  |  |
| 5 | All standard accessories needed to run the bone mill should be included (cables, sterilization container and any other needed accessories) |  |  |
| 6 | All consumables (blades with different sizes and any other consumables) to be quoted separately |  |  |

| Item A-9 | Irrigation-Coupled Radiofrequency Bipolar Tissue Sealer | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Radiofrequency bipolar tissue sealing machine coupled with controlled saline irrigation system dedicated for spine neurosurgeries to provide tissue sealing at low temperatures (~100ºC) of soft tissue and bone with less bleeding, smoke or char compared to conventional electrosurgical units. |  |  |
| 2 | Bipolar cutting and coagulation modes |  |  |
| 3 | Adjustable power level and saline flow rate |  |  |
| 4 | Dual pedal footswitch to activate cut/ coagulation |  |  |
| 5 | Digital LCD display of actual and set power level and irrigation flow rate |  |  |
| 6 | Audible and visual alarm with self-test capability |  |  |
| 7 | Maximum Cut output power ≥ 100 watt, Coagulation output power ≥ 75 watt |  |  |
| 8 | Leakage current ≤ 20 mA |  |  |
| 9 | Small in size and should be able to be mounted on surgical pendant shelf without causing any interferences and signal distortion to nearby equipment |  |  |
| 10 | Safety type : CF defibrillation proof |  |  |
| 11 | All accessories and consumables needed for orthopaedic-spine surgeries should be listed and priced separately, prices should be fixed as per special terms. |  |  |

B:  Vascular

| Item B-1 | 3D C-Arm Imaging System, Vascular | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | A state of the art digital Mobile three dimensional C-Arm fluoroscopy system is required to provide excellent image quality and minimal dose in vascular surgeries. The system must comprise sufficiently sophisticated hardware and software platforms to be able to host a wide range of real-time and post processing image enhancement techniques, the system must meet or transcend the following specifications: |  |  |
| 1 | Motorized lateral and orbital rotation in 3 dimensions (Coaxial, Coronal & Sagittal) for spinal and neurosurgical interventions |  |  |
| 2 | ## Physical Configuration: Two wheeled units with smooth manual steering system: one supporting the C-Arm and console and the other supporting monitors, image processing, recording devices etc. |  |  |
| 3 | ## System backbone: Microprocessor control over all subsystems with an elaborate self-test and error code scheme. |  |  |
| 4 | Control Console: hygienic touch type controls with easy to read, user friendly annotation. |  |  |
| 5 | X-Ray Tube: |  |  |
|  | - Rotating anode type. |  |  |
|  | - Anode heat Capacity, kHU ≥ 300 KHU |  |  |
|  | - Anode cooling, kHU/min ≈ 50 kHU/min |  |  |
|  | - Dual focus, size: ≤ 0.6 each |  |  |
| 6 | X-Ray Generator: |  |  |
|  | - Ripple-free high-frequency converter type. |  |  |
|  | - Power rating, kW@100 kVp: ≥ 15 kW. |  |  |
|  | - Radiographic Mode: - kV range: 40 to 120 kV or better in steps of 1 kV. - mAs range: up to 60 mA or better - Automatic exposure control |  |  |
|  | - Fluoroscopic Mode: - kV range: 40 to 110 kV or better - mA range: Up to 8 mA or better - Pulsed fluoroscopy: 1 to 25 pulses per second or better - Digital snapshot mode |  |  |
| 7 | Flat panel detector |  |  |
|  | - Size not less than (30 cm × 30 cm). |  |  |
|  | - Pixel size ≤ 200 µm or less |  |  |
|  | - CMOS flat panel detector type |  |  |
|  | - DQE ≥ 70% |  |  |
| 8 | TV Monitor: |  |  |
|  | - Dual high resolution displays of at least 18 inches coloured LCD monitors, or one 32 inches coloured LCD monitor |  |  |
|  | - Mounted on an articulating arm. |  |  |
| 9 | Image Acquisition: |  |  |
|  | - Image matrix: ~ 1024x1024. |  |  |
|  | - Collimation: Dual Leaf or Iris type. |  |  |
|  | - Image reversal and rotation capability. |  |  |
|  | - Magnification (detector zoom) |  |  |
| 10 | Image Processing and Storage: |  |  |
|  | - Image storage matrix: ~ 1024x1024. |  |  |
|  | - Patient data registration system. |  |  |
|  | - Last image hold. |  |  |
|  | - Cine replay. |  |  |
|  | - Digital subtraction. |  |  |
|  | - CD/DVD R archiving. |  |  |
|  | - Full DICOM 3.0 connectivity. |  |  |
| 11 | C-Arm: |  |  |
|  | - Free Space: 70cm |  |  |
|  | - Depth: 60cm |  |  |
|  | - Vertical Travel: 40cm |  |  |
|  | - Horizontal Travel: 20cm |  |  |
|  | - Orbital rotation ≥ 120˚ |  |  |
|  | - Pivot rotation ≥ ± 200˚ |  |  |
|  | - Panning motion ≥ 10˚ |  |  |
| 12 | A list of standard accessories for the offered model (optional). |  |  |
| 13 | Offer must include the price for a comprehensive service contract (including all replacement parts and x-ray tubes) valid for a period of (8) years starting at the end of the 24 month warranty period |  |  |

| Item B-2 | Diagnostic Ultrasound System Dedicated for Vascular Applications, Med Range | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | Diagnostic ultrasound system dedicated for vascular applications. The system must meet or transcend the following specifications: |  |  |
| A | Application: |  |  |
| 1 | Abdominal, Musculoskeletal, Neonatal Pediatric, Small Parts, Vascular & Transcranial |  |  |
| B | ## Configuration |  |  |
| 1 | ## Heavy duty Mobile cart / built-in trolley type |  |  |
| 2 | Four full swivel castors with brakes |  |  |
| 3 | Full control keyboard |  |  |
| 4 | High-resolution LCD: ≥ 21” at least |  |  |
| 5 | Storage for transducer cables |  |  |
| 6 | Auxiliary storage compartments/drawers |  |  |
| 7 | Minimum of 4 active transducer connector |  |  |
| 8 | User-adjustable presets |  |  |
| 9 | Adjustable monitor position: up/down, rotate, swivel, & tilt |  |  |
| C | General Features: |  |  |
| 1 | Gray-scale levels: 256 |  |  |
| 2 | Power up time: ≤ 2 minutes |  |  |
| 3 | Self-diagnostic routine during power up |  |  |
| 4 | Acquisition frame rate: ≥ 1000 fps |  |  |
| 5 | Simultaneous processing bandwidth capability: 1-15MHz |  |  |
| 6 | No. of digitally processed channels: ≥ 1000000 |  |  |
| 7 | Input dynamic range: : ≥ 280 dB |  |  |
| 8 | Cine loop: ≥ 1200 frames |  |  |
| 9 | Automatic real-time Doppler trace capability including calculation and display of user-selected measurements |  |  |
| 10 | Measurements on replay |  |  |
| 11 | Real-time Pan/Zoom |  |  |
| 12 | Frozen image Pan/Zoom |  |  |
| 13 | Automatic optimization of base line and scale |  |  |
| 14 | Local HDD patient search through name and/or ID |  |  |
| 15 | Full DICOM 3.0, RIS, HIS, PACS Support |  |  |
| 16 | Raw data processing capability or equivalent |  |  |
| 17 | System must be upgradable, hardware and software |  |  |
| D | Image Storage: |  |  |
| 1 | Capacity of stored images HDD: ≥ 300 GB |  |  |
| 2 | Removable Storage: CD/DVD, USB |  |  |
| E | Multi-frequency Transducers: |  |  |
| 1 | Matrix Linear Array: Please specify the frequency |  |  |
| 2 | Curved Array: 2-6 MHz approximately |  |  |
| 3 | Linear Array : 3-8 MHz |  |  |
| F | Imaging Modes: |  |  |
| 1 | 2D |  |  |
| 2 | B-mode |  |  |
| 3 | Tissue harmonic imaging |  |  |
| 4 | Non-Doppler flow imaging |  |  |
| 5 | M-mode and colour M-mode |  |  |
| 6 | Dual imaging |  |  |
| 7 | Doppler imaging |  |  |
| 8 | Colour Doppler imaging |  |  |
| 9 | Colour Power Angio Imaging |  |  |
| 10 | Pulsed Wave (PW) Doppler |  |  |
| 11 | Continuous wave (CW) doppler |  |  |
| 12 | Simultaneous PW Doppler and 2D (Dual Mode) |  |  |
| 13 | Simultaneous 2D, colour Doppler, and PW Doppler (Triple Mode) |  |  |
| 14 | Compounding imaging |  |  |
| 15 | Speckle reduction ( Noise reduction imaging) |  |  |
| G | Online UPS with backup battery for at least 20 minutes of system operation |  |  |
| H | All available standard & optional features, packages, & accessories must be listed and priced separately |  |  |

SPECIAL TERMS

- Offers not complying with any of the special terms or the technical specifications shall be considered non-conforming with tender requirements.

- Any vendor providing FORGED documents shall be disqualified from the current tender and banned from participating in any future RMS tenders.

- Wherever term “based” is mentioned it refers to the country where the manufacturing company is founded & established.

- All equipment must be the most recently released model/version which is equal to or higher than the range of the specifications of the required system (low, mid or high) and equal to or higher than the level of technology and required options mentioned in  the technical specifications.

- Required certificates (must be submitted with the technical offer):

- FDA clearance (510K) for equipment of USA based companies.

- MDR (EU) 2017/745 certificate for equipment of EU (European Union) based companies.

- UKCA certificate For Equipment of Great Britain based companies (England, Scotland, Northern Ireland, and Wales).

- ARTG (Australian Register of Therapeutic Goods) certification or approval for Australian and New Zealand based companies.

- Establishment License with the relevant Device License issued by the Therapeutic Products Directorate for Equipment of Canadian based companies.

- PMDA (Pharmaceuticals and Medical Devices Agency) certification or approval for equipment of Japanese based companies.

- Swissmedic (Swiss agency for therapeutic products) certification or approval for equipment of Swiss based companies.

- Norwegian Medicines Agency certification or approval for equipment of Norwegian based companies.

- Only for class Ι medical equipment manufactured by companies based in one of the countries mentioned above, submission of either one of the certificates mentioned above or a free sale certificate in any of these countries shall be accepted.

- For the following equipment:

- All X-ray equipment, MRI, ultrasound, and nuclear medicine systems (regardless where the manufacturing companies are based).

- Where the manufacturing companies are based in other origins than the mentioned in terms 2.1 - 2.8.

The following are required:

- At least two of the certificates mentioned above, one of which has to be FDA clearance (510K) (Only for class Ι medical equipment submission of certificates mentioned in 2.9 shall be accepted).

- Evaluation certificate from the Royal Medical Services for the same offered model with at least 80% passing grade.

If the evaluation is not applicable (based on purchasing committee perspective) bidder should submit a list of installation basis of the same offered model and/ or previous models in at least two of the following hospitals (King Hussein Cancer Center, National Center for Diabetes Endocrinology and Genetic Diseases, Jordan University Hospital or King Abdullah University Hospital) with at least three years of operation, list should include: Name of hospital, Model installed, Quantity, and date of installation.

The purchasing committee has the right to officially contact any of these hospitals and disqualify any offer where the feedback is negative in operation, after sales service or local agent performance.

- The vendor is responsible to ensure through official  documents that classified medical devices are manufactured in conformity with applicable quality system standards (ISO, IEC); (the international quality systems standard for medical device manufacturing ISO 13485) or ISO9001.

- With each offer, bidders must provide a formally endorsed document issued by the manufacturer stating that the bidder is the sole certified agent for the offered item.

- In all of the above cases (except 2.12) certificates must be formally endorsed by JFDA.

- Any vendor not submitting all required certificates will be eliminated.

- Official website of the manufacturer must be provided. The website must clearly demonstrate the history profile and manufacturing features of the company including the same offered model and its brochure.

In case of inconsistency between the information on the website with the provided information of the local agent regarding any offered product, the committee has the right to eliminate that product from the awarding process.

- Offered items should be from reputable well known manufacturers and excellent experience in the field and shall have multiple installations of the same offered model and/ or previous models in RMS main hospitals with at least two years of operation and excellent experience in operation, after sales service & local agent performance; otherwise the purchasing committee has the right to request any of the following:

- An evaluation certificate as mentioned in term 2.10.b.

- A sample of the same offered model at any time during the purchasing process, the sample unit shall be delivered within three weeks from the date of a written notification; any vendor rejects to deliver a sample unit will be eliminated. Evaluation duration will be determined by the purchasing committee.

Any offered item fail in the evaluation/assessment process will be rejected

- Vendors must specify the origin of the offered items and accessories in the technical offer.

- Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first half of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period shall be rejected.

- Goods are to be brand new; manufacturing date of the delivered products should not exceed 18 months from the date of the final award.

- Delivery period should be mentioned clearly; the purchasing committee has the right to reject any offer with delivery time exceeding 4 months.

- Warranty:

- Offers must include a full warranty for a period of a minimum of 24 months from the date of installation or 30 months from the date of receiving the items at the agreed location mentioned in the final order (whichever comes first).

Warranty must include corrective & preventive maintenance activities as per manufacturer recommendations including:

- Required spare parts (free of charge)

- Labour

- Hardware

- Software

- Rechargeable batteries

At the end of the warranty period, The supplier commits to implement final inspection of the submitted goods and submit the reports signed by the site chief engineer stating that the equipment are working properly as well as all preventive maintenance reports during the warranty period.

Warranty & installation are excluded for items mentioned in Attachment no.1 the purchasing committee has the right to exclude any other item not listed in the attachment and does not required installation and warranty.

- In the case where a delay in installation has occurred as a result of the supplier's dereliction and has exceeded a period of one month from the date of receiving the items, the approved warranty period will automatically be considered as 24 months from the date of installation.

- If at any time during the warranty period the item becomes inoperative due to a technical fault the item must then be repaired by the supplier /local agent within a period of 5 working days from written notification, warranty will be extended according to downtime period.

- If the delay exceeds 30 days the supplier must replace the item with a new identical functioning one (within the same delivery period mentioned in the final offer). In case the item was replaced by a new one, the warranty period mentioned in 9.a) above will start from the installation and commissioning date of the new item.

- Local agent/ supplier is committed to transport and install any awarded item during the warranty period on free of charge basis to any location inside the country whenever required by Royal Medical Services, this should not include pre-installation requirements in the new site.

- All offered items (main unit) should be fully designed, manufactured, and labelled by their real original manufacturer in which all related testing, research, development and approvals went through.

Any relabelled products for the main unit (white-label manufacturing, OEM, or repackaging) are rejected.

- One set of operation manual(s) and one set of service manual(s) including schematics and a spare-part list must be delivered with each unit, CD/DVD is acceptable. For large tenders, a certain agreed percentage of manuals per item may be agreed upon.

- Where applicable, pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include removal of old system(s), any civil work, electrical work or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new system(s).

- Power requirements: where applicable either single phase 220V, 50Hz or 3-phase 380V. Systems with external transformers are considered conforming only if clearly stated in the technical specifications.

- Technical offers must include clear original technical brochures/catalogues for all offered items.

- Offers must include fully detailed technical offers and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s), full technical description/specifications and any accessories or options included in the offer.

Offer must clearly indicate the origin of the offered item and the country where the manufacturer of the offered items is based.

- Compliance sheets must be as per the tabular format of the technical specifications in the tender documents, listing the required specifications on one column and a  Yes  or NO  response to each point in the adjacent column, with reference to page and line numbers in the relevant technical brochure. Offers not complying with this term shall be rejected.

- Qualifications and after sales service: The technical bid should contain all the necessary documents (training certificates) to prove the capability of the bidders for supplying and installing a trouble free equipment meeting the quality standards and technical specifications of the manufacturer and the ability of the bidders for providing efficient after sales service conducted by authorized certified biomedical service engineers with minimum 2 years experience in the same field.

- Accessories and consumables:

  - Any accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately.

All offered accessories and consumables must be approved by the manufacturer.

  - Technical offers must include a priced list for all accessories and consumables related to the required equipment as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) with prices fixed for a period of five years from the date of installation and commissioning with a maximum annual increase of 2%, any essential item not listed will be considered free of charge.

  - Accessories and consumables must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Where applicable, a start-up kit of accessories and consumable items must be provided with each system on a free-of-charge basis.

- Spare Parts:

  - Technical offers must include a comprehensive and priced list of all spare parts related to the awarded equipment (including rechargeable batteries) as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) valid for a minimum period of five years with a maximum annual increase of 2%, commencing at the end date of the warranty period, any essential item not listed will be considered free of charge.

  - Spare parts must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Prices of spare parts should be reasonable and will be taken into consideration during the purchasing process; the purchasing committee has the right to eliminate any offered item with unreasonable high prices of spare parts.

  - Delivery period of required spare parts should not exceed 2 months from the date of the final order.

- Spare parts, consumables and accessories availability must be guaranteed for a minimum period of ten years starting from the date of installation and commissioning.

- Local agent is committed to sign a service contract with Royal Medical Services to perform all relevant maintenance for the awarded equipment (whenever required by DRMS within a period that does not exceed one year from the end date of the warranty period and if mentioned clearly in the specifications) without exceeding the original mentioned cost and as per mentioned terms within minimum required period.

- Tender Awards:

  - For the final list of offers having a chance of winning the award, the awarding process shall be based on the accumulative value of both the offered item and its’ running cost (Total Cost of Ownership) over a period of seven years from the date of installation and commissioning. Only offers with the lowest total cost of ownership over a period of seven years from the date of installation and commissioning shall qualify for the award.

  - Running cost includes the value of consumables, accessories needed to operate the system over the same period as well as the cost of any service contract (where applicable).

- For PC/Laptop based systems:

  - Complete restoration medium (CD/DVD/etc.) of the operating system and the application software must be supplied.

  - Purchased licenses, software keys & dongles must be provided at any time during and after the warranty period on free of charge basis whenever required by DRMS for a minimum period of ten years starting from the date of installation and commissioning.

  - Where locally supplied computers, laptops & printers are offered, the offered model should be from well known manufacturer.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores, installation, pre-installation (if needed), training, commissioning, warranty and bringing the equipment into service.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- a. DRMS has the right to increase the awarded quantities by a percentage not exceeding 35% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

b. DRMS has the right to decrease the awarded quantities by a percentage not exceeding 50% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- a. Bidder is not allowed to submit more than one offer for the same item whether that’s solely or in coalition or partnership with other bidders.

b. Bidder is allowed to include within their offer optional alternatives for the same offered item from the same manufacturer.

- The supplier must furnish DRMS with a guarantee stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelve months from the date of final acceptance of the equipment by DRMS.

- Training:

  - For items where abroad service training courses are required in technical specifications, offers must include a certified service training program for at least 3 working days at a reputable center abroad recognized by the manufacturer for at least one biomedical engineer or biomedical technician; all costs inclusive, air tickets, , boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - For items where abroad user  training courses are required in technical specifications, offers must include a certified operator training program at a reputable center abroad recognized by the manufacturer for at least one operator; all costs inclusive, air tickets, boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - The period of the training courses must be according to the manufacturer’s program excluding traveling days and must be stated clearly in the technical offer.

  - Training Programs must conform to the following standards:

- User training must comprise understanding and use of operation manual(s), correct and safe operation of the equipment, as well as user preventive maintenance and calibration.

- Service training must comprise: theory, understanding and use of service manual(s), calibration, preventive maintenance procedure, and practical troubleshooting and repair exercises, and must be conducted by professional instructors employed or authorized by the system manufacturer.

- Service training must be conducted on a system of identical make, model, and configuration to that purchased by DRMS, and designated by the manufacturer or the local agent for training purposes.

- Certificates must be endorsed and officially sealed by the system manufacturer, legally empowering trainees to engage in user and service activities according to operation and service manual(s).

- Where applicable, offers must include an on-site user and service training.

30. For offers submitted in Jordanian dinars, Payment will be either by wire transfer or by cheque after final acceptance of goods. Any other way of payment will be rejected.

Attachment 1

| # | Equipment |
| --- | --- |
| 1 | Adjustable Auto Pipettes |
| 2 | Resuscitation Bag |
| 3 | Laryngoscope Set |
| 4 | Oxygen Flow meter wall type/ single |
| 5 | Regulator Suction with canister, wall vacuum outlet |
| 6 | Oxygen Regulator for Oxygen Cylinders |
| 7 | Pulse Oximeter, Finger type |
| 8 | Oxygen Cylinder |
| 9 | Doppler, portable |
| 10 | Diagnostic set, Portable |
| 11 | Direct Ophthalmoscope, Portable |
| 12 | Otoscope, Portable |
| 13 | Air Mattress System, homecare |
| 14 | Stethoscope |
| 15 | Aneroid Sphygmomanometer |
| 16 | Video Assisted Laryngoscope, portable |
| 17 | Scale, Manual |
| 18 | Wood's Light |
| 19 | Cough Pressure, Normal Saline |
| 20 | Rehabilitation Walking Parallel Bars, non-powered |
| 21 | Therapy Mat |
| 22 | Medical Ball All Size |
| 23 | Dumbbells Rack with complete set of dumbbells |
| 24 | Crutches |
| 25 | Shoulder wheel |
| 26 | Mobile Mirror |
| 27 | Cuff Weights |
| 28 | Walker, different sizes |
| 29 | Patient Elbow Stick |
| 30 | Rehabilitation Training Ladder |
| 31 | Rehabilitation Suspension Frame |
| 32 | Exercise Band All Size (Theraband) |
| 33 | Lens trial set |
| 34 | Wheel Chair |
| 35 | Commode Chair |
| 36 | Bassinet (Baby Cot) |
| 37 | Resuscitation Cart (Crash Cart) |
| 38 | Medication Cart |
| 39 | Cart, Drawers |
| 40 | Examination Couch, Manual |
| 41 | Gynaecology Examination Table, Manual |
| 42 | Examination Table, Neonates, Manual |
| 43 | Intravenous Pole, Mobile Stand |
| 44 | DDA Cabinet |
| 45 | Stainless steel Multipurpose Trolley |
| 46 | Cabinet, Instrument, Operation Theatres |
| 47 | Dressing Cart |
| 48 | Stainless steel wire shelving unit |
| 49 | Paper Trolley |
| 50 | Stainless Steel Sink (Clean up counter) |
| 51 | Scopes Cabinet |
| 52 | Mayo Table |
| 53 | Table, Instrument |
| 54 | Stool, Adjustable, Doctor |
| 55 | Stool, Adjustable, Operation Theatres |
| 56 | Carts, linen/laundry, soiled, Double |
| 57 | Closed distribution trolley |
| 58 | Step Ladder, Conductive, Double |
| 59 | Step, Surgeon, Single |
| 60 | Kick Bucket |
| 61 | Mobile Stand for Oxygen Cylinder |
| 62 | Stainless Steel Wire Basket, 1 STU |
| 63 | Cart, Plaster |

Annex (4) ENT $ Dental

A: ENT

| Item A-1 | ENT Surgical Microscope | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # A floor mounted dedicated surgical microscope for use in ENT surgeries. |  |  |
| 2 | # Magnification system that features motorized zoom with 1:6 ratio. |  |  |
| 3 | # Binocular Tube : |  |  |
|  | - Two binocular tubes with apochromatic eyepieces for (Main, and assistance) # Binocular assistant microscope should be synchronized with the main one with the same light route, magnification viewing field and zero degree |  |  |
|  | - Tiltable tubes 0˚ - 180˚ |  |  |
|  | - Main eyepieces with at least 12.5 x magnification |  |  |
|  | - Assistant eyepieces with at least 10 x magnification |  |  |
| 4 | # Working distance with a range of 225- 450 mm or wider |  |  |
| 5 | # Auto focus |  |  |
| 6 | # Multifunctional 2 handles with programmable buttons |  |  |
| 7 | # X-Y movement ability: Microscope should include X-Y coupling for fine adjustment of microscope head |  |  |
| 8 | # Integrated high-resolution camera, HD or better |  |  |
| 9 | # Arm to reach at least 1300 mm |  |  |
| 10 | # At least 21" Medical grade touch LCD monitor |  |  |
| 11 | # Electromagnetic braking system |  |  |
| 12 | # Xenon cold light illumination system that provides minimum of 300 watt with integrated emergency lamp (not halogen) # Replacing lamps and modules should be easy and simple. |  |  |
| 13 | # Original digital video system with archiving capability on DVD or USB memory that is supplied and approved by the same manufacturer of the main system |  |  |
| 14 | # Should be compatible with HAKEEM hospital information system, all hardware, software and licenses should be included |  |  |
| 15 | # Foot brake for locking of moveable castors |  |  |
| 16 | # Switches and controls should be sealed against penetration of liquids |  |  |
| 17 | # Original Advanced ENT Surgeon chair with following specs: - Motorized up and down - Rechargeable battery with charger - High stability brake system |  |  |
| 18 | # Any other additional features available in the system are to be mentioned for comparison purposes. |  |  |
| 19 | # All accessories and options must be listed and priced separately. |  |  |
| 20 | # A list of standard accessories for the offered model (optional) |  |  |
| 21 | # Abroad service training for at least 1 biomedical engineer/ technician as per special terms |  |  |

| Item A-2 | ENT Sinuscopy Surgical Tower | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| A | A top of the line diagnostic endoscope for an ENT for operating room is required to have the following minimum requirements |  |  |
| B | Camera system (4k Resolution) : Provides electronic shutter function, automatic brightness control and optical zoom |  |  |
|  | - Resolution (4k): a minimum of 3840*2160. |  |  |
|  | - Compact, lightweight, easy to grip camera head with low heat generation |  |  |
|  | - Focal Distance with zoom lens, 16-28 mm or better |  |  |
|  | - Gas sterilizable, autoclavable is preferable |  |  |
|  | - Immersible in disinfections liquid. |  |  |
|  | - Universal coupling mechanism to fit all standard endoscopes. |  |  |
|  | - Camera cable 3 m or more |  |  |
|  | - Camera controller (4k resolution) |  |  |
|  | - Picture in picture mode |  |  |
| C | Light Source (either standalone device or integrated on camera control unit): |  |  |
|  | - LED Light source (light intensity should be equivalent to 175 Watt xenon lamp) |  |  |
|  | - Working hour ≥ 30,000 hours. |  |  |
|  | - Automatic and manual adjustment of light intensity. |  |  |
|  | - Front panel hygienic with convenient indicators for ease of operation. |  |  |
|  | - Flexible fibre optic cable preferably autoclavable and not less than 230 mm. |  |  |
| D | LCD / LED Colour Monitor: |  |  |
|  | - Medical grade type that is supplied by the same manufacturer |  |  |
|  | - At least of 32 inch size that is mounted on the mobile cart |  |  |
|  | - Resolution: 4k resolution minimum of 3840*2160 |  |  |
|  | - DVI compatible. |  |  |
| E | Ability to record images and videos on Medical grade DVD recorder or USB flash memory |  |  |
| F | Rigid nasal telescopes (4k Resolution), sterilizable with protection cover: |  |  |
|  | - Degree (0), Diameter (4.00 mm), working length ≥ 15cm Qty. (2) |  |  |
|  | - Degree (30), Diameter (4.00 mm), working length ≥ 15cm Qty. (2) |  |  |
|  | - Degree (30), Diameter (4.00 mm), working length ≥ 15cm Qty. (2) |  |  |
|  | - Degree (70-90), Diameter (4.00 mm), working length ≥ 15cm Qty. (1) |  |  |
| G | Rigid ear telescope (4k Resolution) with the following specifications: |  |  |
|  | - Zero degree |  |  |
|  | - 4 mm inner diameter |  |  |
|  | - Working length approx. 6cm |  |  |
| H | ICG option should be available and quoted separately, 4K Camera, Camera controller, Light source & telescopes should be ICG compatible |  |  |
| I | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for the subcomponents for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |

| Item A-3 | High speed Otologic Surgical Drill | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Specially designed for ENT surgery |  |  |
| 3 | Main Control Drive |  |  |
|  | - Can fit all hand-pieces |  |  |
|  | - The unit must include two outlets for different saws and drills or better |  |  |
|  | - Capability to recognize hand-pieces |  |  |
|  | - Capability to control speed and directions |  |  |
|  | - Integrated irrigation pump |  |  |
|  | - Filters to insure sterility |  |  |
|  | - Footswitch must be included |  |  |
| 4 | Electric Micro drill handpiece Qty. (2) |  |  |
|  | - High Speed otologic drill 60,000 rpm or better operates nasal &laryngeal in forward and oscillating way |  |  |
|  | - Straight attachment Qty. (2) |  |  |
|  | - Angled attachment Qty. (2) |  |  |
|  | - Different types and sizes burs (priced separately) |  |  |
|  | - Irrigation system must be included at integrated console pump |  |  |
|  | - Speed and irrigation must be controlled from foot switch and main control drive |  |  |
|  | - Irrigation cassette and clip must be included (priced separately) |  |  |
|  | - Autoclavable 134℃ |  |  |
| 5 | Electric shaver handpiece Qty. (2) |  |  |
|  | - Easy to fit blades without any tools |  |  |
|  | - Speed 5000rpm or better |  |  |
|  | - Different types of blades |  |  |
|  | - Irrigation cassette and clip must be included |  |  |
|  | - Autoclavable 134℃ |  |  |
| 6 | Special drill for middle Ear surgery (Skeeter) or equivalent Qty. (1) |  |  |
| 7 | Sterilization container for all hand pieces and attachments, burs and blades to be included. |  |  |
| 8 | Standard original trolley with all accessories |  |  |
| 9 | Consumable and accessories required for ENT operations must be quoted separately. |  |  |
| 10 | # A list of standard accessories of stainless steel (cutting burs, blade, Skeeter burs ….. etc.), must be offered (reusable or disposable) |  |  |

| Item A-4 | Intra-Operative Nerve Locator/Stimulator, ENT | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The system must be the latest model that is suitable for facial nerve monitoring |  |  |
| 2 | 4- Channels EMG monitoring capability |  |  |
| 3 | # Audible and visual real-time feedback upon nerve irritation to aid in intraoperative decision making |  |  |
| 4 | # Large coloured touch screen |  |  |
| 5 | # Patient Interface (wired or wireless) between electrodes and the screen |  |  |
| 6 | # Stimulator Probe, monopolar |  |  |
| 7 | # Muting detector probe to reject noise generated from electrosurgical units |  |  |
| 8 | # EMG recording capability |  |  |
| 9 | # Original cart to be included |  |  |
| 10 | # Printer should be quoted separately |  |  |
| 11 | # All consumables including (probes and sub-dermal needle electrodes) should be listed and quoted separately |  |  |
| 12 | # Options to be quoted separately |  |  |
| 13 | # All other accessories to be priced separately |  |  |

| Item A-5 | ENT Head Light, Wireless | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Portable, light weight, and compact design |  |  |
| 2 | Wireless headband with integrated rechargeable battery |  |  |
| 3 | Adjustable light intensity and spot size |  |  |
| 4 | Rapid adjustment of cross-headband |  |  |
| 5 | LED illumination type with light intensity up to 30,000 lux or better |  |  |
| 6 | LED operating life time 30,000 hours or better |  |  |
| 7 | Rechargeable battery operation not less than 3-hrs or better with fully charged battery |  |  |
| 8 | Battery charger device |  |  |
| 9 | Two backup rechargeable batteries to be included and priced separately. |  |  |
| 10 | All accessories or options should be quoted and priced separately |  |  |

| Item A-6 | CO2 Laser System | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # A mobile microprocessor controlled laser system that utilizes carbon dioxide as its laser medium to produce and route for use a beam of energy 10,600 nm to operate in the continuous wave (CW) and pulsatile modes |  |  |
| 2 | # An energy source (a high-voltage DC or RF energy source) to electrically excite the active medium. Featuring the following: |  |  |
|  | # 1-25 watt output adjustable in increments not greater than 5% of the current setting of 1w, whichever is greater. |  |  |
|  | # Accuracy <10% of the setting |  |  |
|  | # The displayed power setting should indicate the power delivered to tissue. |  |  |
| 3 | # A cooling system (Fully built-in, not to require an external water supply |  |  |
| 4 | # A control panel: touch screen |  |  |
| 5 | # The output characteristics (e.g. operation mode, pulse duration, output power, exposure time) should be clearly displayed. |  |  |
| 6 | # Audible and/or visual alert to indicate any made change in output characteristics. |  |  |
| 7 | # Should clearly indicate malfunction of microprocessor or control circuitry. |  |  |
| 8 | # Switches and controls should be protected against accidental setting changes, and should be sealed against penetration of liquids. |  |  |
| 9 | # Exposure duration: |  |  |
|  | # Adjustable |  |  |
|  | # Accuracy of automatically controlled pulse duration should be within 10% of set values. |  |  |
|  | # Pulse counts should not be in error by more than one count or 5% whichever is greater. |  |  |
| 10 | # Aiming laser beam: Adjustable intensity of aiming beam, Bidders shall specify type of aiming beam |  |  |
| 11 | # A delivery system: Through fiber-optic cable using a spring loaded articulated arm (arm length 1300mm approx.). |  |  |
| 12 | # Adaptor and micromanipulator to connect CO2 Laser with the surgical microscope (in item above) |  |  |

| Item A-7 | Intra-Operative Navigation System, ENT | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | An Intra-Operative Navigation System is required from a reputable well known company that has an excellent experience in this field and has multiple installation bases in Jordan, a list of installation bases should be included in the technical offer |  |  |
| 2 | Mobile On trolley with antistatic castors |  |  |
| 3 | Tracking System: - Infrared camera for optical wireless marker tracking |  |  |
| 4 | Transfer and planning: - Platform for advanced surgery planning - DICOM compatible to transfer MRI and CT images to navigation system from the available RIS network, all licenses should be included - Ability to download MRI/ CT images from USB and CD/ DVD - High speed CPU - At least 1 TB HDD - At least one high resolution touch-screen monitors ≥ (27 inches) |  |  |
| 5 | ENT Image registration: - Accurate, simple and fast with less than 2 minutes - Method: Touch by tracing patient’s anatomy (with & without using fiducials) or touchless by laser pointer |  |  |
| 6 | ENT Software: |  |  |
|  | - User friendly operating system |  |  |
|  | - Automatic image fusion software: - Fast and exact image fusion - Supports data of CT, MR (T1, T2 and MRA) |  |  |
|  | - Pre-planning of multiple trajectories (target/ entry point) |  |  |
|  | - Zoom function and adjustable grey level windowing |  |  |
|  | - Overview of all patient images for quick image selection, at least 4 images/ patient |  |  |
|  | - 3D object reconstruction and planning |  |  |
|  | - Fast visualization of volumetric information in axial, coronal and sagittal planes |  |  |
|  | - Visual tool-tip extension for surgery simulation |  |  |
|  | - Storage of screen shots for scientific documentation |  |  |
| 7 | Universal instrument tracking adaptors, hardware & software for the identification of straight Instruments) |  |  |

| Item A-8 | Otoscope, Portable | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, compact, heavy duty construction & reliable, ergonomic design and easy to use |  |  |
| 2 | Power handle & Otoscope head with at least three different sizes of reusable specula. |  |  |
| 3 | A hard-shell well-padded protective case. |  |  |
| 4 | Power handle should be made from heavy duty, easy to clean, corrosion resistant material. |  |  |
| 5 | Batteries and light bulbs should be commonly available from several known manufacturers. |  |  |
| 6 | Rechargeable type batteries |  |  |
| 7 | Battery charger should be included |  |  |
| 8 | LED illumination type |  |  |
| 9 | Kindly quote for all sizes and types of specula / accessories. |  |  |

B: Dental

| Item B-1 | Arthroscopic Surgical Tower for Oral and Maxillofacial Surgeries | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| A | Durable, heavy duty construction, reliable with superior image quality |  |  |
| B | Camera system (4k Resolution): Provides electronic shutter function, automatic brightness control and optical zoom |  |  |
|  | - Resolution (4k): a minimum of 3840*2160. |  |  |
|  | - Compact, lightweight, easy to grip with low heat generation |  |  |
|  | - Focal Distance with zoom lens, 16-28 mm or better |  |  |
|  | - Gas sterilizable, autoclavable is preferable |  |  |
|  | - Immersible in disinfections liquid. |  |  |
|  | - Universal coupling mechanism to fit all standard endoscopes. |  |  |
|  | - Camera cable 3 m or more |  |  |
|  | - Camera controller (4k resolution) |  |  |
|  | - Picture in picture mode |  |  |
| C | Light Source (either standalone device or integrated on camera control unit): |  |  |
|  | - LED Light source (light intensity should be equivalent to at least 170 Watt xenon lamp) |  |  |
|  | - Working hour ≥ 30,000 hours. |  |  |
|  | - Automatic and manual adjustment of light intensity. |  |  |
|  | - Front panel hygienic with convenient indicators for ease of operation. |  |  |
|  | - Flexible fibre optic cable preferably autoclavable and not less than 230 mm. |  |  |
| D | Inflow/outflow pump: |  |  |
|  | - Touch screen ensure safe and precise adjustment of the set values |  |  |
|  | - Parameters display, easy to read and arranged clearly parallel to one another allow the user to monitor the current actual and set values of all unit parameters at any time |  |  |
|  | - audible alarms in case of malfunction |  |  |
|  | - Irrigation Pressure: adjustable up to 150 mmHg or better |  |  |
|  | - Irrigation Flow Rate: adjustable up to 0.5 L/M or better |  |  |
| E | LCD / LED Colour Monitor: |  |  |
|  | - Medical grade type that is supplied by the same manufacturer |  |  |
|  | - At least of 32 inch size that is mounted on the mobile cart |  |  |
|  | - Resolution: 4k resolution minimum of 3840*2160 |  |  |
|  | - DVI compatible. |  |  |
| F | Ability to record images and videos on Medical grade DVD recorder or USB flash memory |  |  |
| G | Motors and shavers |  |  |
| G-1 | Motor system control unit : |  |  |
|  | - Speed range : up to 10000 rpm or better |  |  |
|  | - Dual pedals footswitch |  |  |
| G-2 | Shaver hand piece: |  |  |
|  | - Two modes of operation: Normal and Oscillation. |  |  |
|  | - Adjustable rotational speed up to 7,000 rpm or better |  |  |
|  | - Adjustable oscillation speed up to 3,000 rpm or better |  |  |
|  | - The speed and mode of the shaver hand piece may be controlled by the control unit, the hand piece, or the footswitch. |  |  |
|  | - Fast and simple blade change using a precise, quick-release coupling. |  |  |
|  | - Autoclavable |  |  |
| G-3 | Miniature Shaver Hand piece for small joints (TMJ): |  |  |
|  | - Ø for Shaver tools : 2 - 4.5 mm |  |  |
|  | - Rotational speed: up to 3000 rpm or better |  |  |
|  | - Autoclavable |  |  |
| G-4 | Arthroscopic Drill for small joints (TMJ) to fit with the above console: |  |  |
|  | - Power drill handle rotational speed: adjustable up to 1,000 rpm or better |  |  |
|  | - Universal connecting cable |  |  |
|  | - Jacobs chuck keyed up to 3.9 approximately |  |  |
|  | - Jacobs chuck keyed up to 6.3 mm approximately |  |  |
|  | - Sagittal saw attachment |  |  |
| G-5 | All burrs , resectors , Gide wires, cannulated drill and saw blade sagittal must be quoted separately |  |  |
| H | Instruments: |  |  |
|  | - TMJ Level I Arthroscopy: - Arthroscope Telescope, 30º, Diameter ≈ 1.9 mm, Autoclavable, 4K resolution - Warburton Collar-lock Scope Cannula (≈ 50 mm) - Trocar, ≈1.9 mm - Obturator, ≈1.9 mm - Suction Tube (5 French) - Arthroscopy System Tray |  |  |
|  | - Level II & Level III Arthroscopy - Scope/Operative Cannula, ≈ 1.9 mm - Trocar, ≈ 1.9 mm - Obturator , ≈ 1.9 mm - Probe, Straight, Blunt, ≈ 1.8 mm - Probe, Curved, Blunt, ≈ 1.8 mm - Banana Knife, ≈ 1.8 mm - Scissors, Micro-Style, ≈ 1.8 mm - Laser Fiber Guide, ≈ 1.8 mm - Biopsy Forceps, Serrated, ≈ 1.8mm - Biopsy Forceps, Micro-Style, Serrated, ≈ 1.8mm - Tissue Grasping Forceps, Micro-Style, Serrated, ≈ 1.8mm - Bone Curette, ≈ 1.8mm - Bone File, Concave, Pull, ≈ 1.8mm - Bone File, Convex, Push, ≈ 1.8mm - Golden Retriever, ≈ 1.8mm - Biopsy Forceps, Ergo-Style, Sharp-Cup, ≈1.8mm - Arthroscopy System Tray |  |  |
| I | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for the subcomponents for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |

| Item B-2 | Sialendoscope | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Semi-flexible endoscope for exploring the salivary ducts and removing salivary stones |  |  |
| 2 | Integrated irrigation and working channels |  |  |
| 3 | 5° curvature at the distal tip |  |  |
| 4 | Enables the minimally invasive removal of obstructions in the salivary ducts |  |  |
| 5 | Scope and accessories should be autoclavable |  |  |
| 6 | Accessories: the following should be priced separately: - Dilator for salivary duct, diameter 0.5-0.8 mm, length 11 cm, approximately - Dilator for salivary duct, diameter 0.7-1.0 mm, length 11 cm, approximately - Dilator for salivary duct, diameter 0.9-1.2 mm, length 11 cm, approximately - Dilator for salivary duct, diameter 1.0-1.6 mm, length 11 cm, approximately - Dilator for salivary duct, diameter 1.3-1.8 mm, length 11 cm, approximately - Dilator for salivary duct, length 14 cm, approximately - Dilator for Salivary Duct, pointed, for Wharton duct - Micro burr for salivary stones, diameter 0.35 mm approximately - Cleaning Brush, round, flexible, for working channel diameter 0.3 - 0.4 mm, length 40 cm, approximately - Cleaning Brush, round, flexible, for working channel diameter 0.6 - 0.8 mm, length 40 cm, approximately - Miniature Straight Forward Telescope 0°, 1.6 mm approximately, semi-flexible, autoclavable, working length 10 cm approximately, irrigation channel 0.25 mm approximately, length 140 cm approximately - Light cable, diameter ≤ 2.5 mm, length ≥ 230 mm - Grasping Forceps, double jaws, flexible, diameter 0,8 mm, working length 30 cm, approximately - Biopsy Forceps, double jaws, flexible, diameter 0.8 mm, working length 30 cm, approximately - Salivary Probes sizes 0, 1, 2, 3, 4, 5, 6, 7 - Sterilization tray for scope and accessories |  |  |

| Item B-3 | Dental Surgical Drill and Saw Unit | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Specifically designed for Oral, Maxillofacial surgeries. |  |  |
| 2 | The system shall include at least the following: |  |  |
| 2.a | Console: |  |  |
|  | - Flat touch control panel to control speed of all drills and saws. |  |  |
|  | - To include irrigation pump either built in or as an external unit. |  |  |
|  | - The unit must include multiple outlets for different saws and drills. |  |  |
| 2.b | Electric Cable of at least 3 meters long to connect console with the hand pieces. |  |  |
| 2.c | Foot switch: |  |  |
|  | - Wired model |  |  |
|  | - To activate micro drill in both directions; forward and reverse. |  |  |
|  | - To activate all saw hand pieces; sagittal, oscillating and reciprocating. |  |  |
| 2.d | Micro Drill Hand piece: |  |  |
|  | - Speed approximately 30,000 rpm or above |  |  |
|  | - Electrically controlled |  |  |
|  | - Must include angled attachments, please quote for different sizes. |  |  |
|  | - Autoclavable at 134° C |  |  |
|  | - Different sizes of burs and drill bits must be quoted separately. |  |  |
| 2.e | Saw hand pieces: |  |  |
|  | - The offer must include sagittal, oscillating and reciprocating saw handpieces. |  |  |
|  | - Electrically controlled hand pieces |  |  |
|  | - Autoclavable at 134° C |  |  |
|  | - The offer must include different sizes of sagittal, oscillating and reciprocating cutting blades quoted separately. |  |  |
| 2.f | Sterilization case |  |  |
| 2.g | Oil spray if needed |  |  |

| Item B-4 | Dental Laser Unit (Soft Tissue) | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The unit must be suitable for soft tissues procedures in dental applications. |  |  |
| 2 | Table-top or portable type compact in design. |  |  |
| 3 | Laser type: Diode laser |  |  |
| 4 | Wavelength: 940 nm approximately |  |  |
| 5 | Low cutting power of less than 5 watt. |  |  |
| 6 | Built in rechargeable batteries |  |  |
| 7 | External AC/ DC adapter/charger is acceptable |  |  |
| 8 | Pencil type hand-piece |  |  |
| 9 | Hand-piece should be autoclavable |  |  |
| 10 | Hand-piece activation control |  |  |
| 11 | All other accessories and consumables (tips, safety goggles, etc...) should be quoted |  |  |

| Item B-5 | Intra-Operative Nerve Locator/Stimulator, Maxillofacial | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The system must be the latest model that is suitable for facial nerve monitoring |  |  |
| 2 | 4- Channels EMG monitoring capability |  |  |
| 3 | # Audible and visual real-time feedback upon nerve irritation to aid in intraoperative decision making |  |  |
| 4 | # Large coloured touch screen |  |  |
| 5 | # Patient Interface (wired or wireless) between electrodes and the screen |  |  |
| 6 | # Stimulator Probe, monopolar |  |  |
| 7 | # Muting detector probe to reject noise generated from electrosurgical units |  |  |
| 8 | # EMG recording capability |  |  |
| 9 | # Small in size and should be able to be mounted on surgical pendant shelf, original cart to be priced separately |  |  |
| 10 | # Printer should be quoted separately |  |  |
| 11 | # All consumables including (probes and sub-dermal needle electrodes) should be listed and quoted separately |  |  |
| 12 | # Options to be quoted separately |  |  |
| 13 | # All other accessories to be priced separately |  |  |

| Item B-6 | Dental Ultrasonic Bone Surgery System | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # Heavy duty, robust from a well-known reputable manufacturer |  |  |
| 2 | # Compact bench top type |  |  |
| 3 | # Operating frequency range: (25-32) KHz approximately |  |  |
| 4 | # Tissue selectivity, the device cut only hard tissue |  |  |
| 5 | # Piezo LED autoclavable hand piece |  |  |
| 6 | # Irrigation pump with all accessories , such as water line, saline bag holder, etc. |  |  |
| 7 | # At least 3 pre-programming mode |  |  |
| 8 | # Screen with touch button to control all parameter |  |  |
| 9 | # Multifunction foot control |  |  |
| 10 | # Original tips Kit, Qty. (2) must be included |  |  |
| 11 | # Cable length ≥1.8 m |  |  |
| 12 | # Sterilization case must be included |  |  |
| 13 | # All accessories needed to operate the device must be included |  |  |
| 14 | All accessories must be priced separately |  |  |

SPECIAL TERMS

- Offers not complying with any of the special terms or the technical specifications shall be considered non-conforming with tender requirements.

- Any vendor providing FORGED documents shall be disqualified from the current tender and banned from participating in any future RMS tenders.

- Wherever term “based” is mentioned it refers to the country where the manufacturing company is founded & established.

- All equipment must be the most recently released model/version which is equal to or higher than the range of the specifications of the required system (low, mid or high) and equal to or higher than the level of technology and required options mentioned in  the technical specifications.

- Required certificates (must be submitted with the technical offer):

- FDA clearance (510K) for equipment of USA based companies.

- MDR (EU) 2017/745 certificate for equipment of EU (European Union) based companies.

- UKCA certificate For Equipment of Great Britain based companies (England, Scotland, Northern Ireland, and Wales).

- ARTG (Australian Register of Therapeutic Goods) certification or approval for Australian and New Zealand based companies.

- Establishment License with the relevant Device License issued by the Therapeutic Products Directorate for Equipment of Canadian based companies.

- PMDA (Pharmaceuticals and Medical Devices Agency) certification or approval for equipment of Japanese based companies.

- Swissmedic (Swiss agency for therapeutic products) certification or approval for equipment of Swiss based companies.

- Norwegian Medicines Agency certification or approval for equipment of Norwegian based companies.

- Only for class Ι medical equipment manufactured by companies based in one of the countries mentioned above, submission of either one of the certificates mentioned above or a free sale certificate in any of these countries shall be accepted.

- For the following equipment:

- All X-ray equipment, MRI, ultrasound, and nuclear medicine systems (regardless where the manufacturing companies are based).

- Where the manufacturing companies are based in other origins than the mentioned in terms 2.1 - 2.8.

The following are required:

- At least two of the certificates mentioned above, one of which has to be FDA clearance (510K) (Only for class Ι medical equipment submission of certificates mentioned in 2.9 shall be accepted).

- Evaluation certificate from the Royal Medical Services for the same offered model with at least 80% passing grade.

If the evaluation is not applicable (based on purchasing committee perspective) bidder should submit a list of installation basis of the same offered model and/ or previous models in at least two of the following hospitals (King Hussein Cancer Center, National Center for Diabetes Endocrinology and Genetic Diseases, Jordan University Hospital or King Abdullah University Hospital) with at least three years of operation, list should include: Name of hospital, Model installed, Quantity, and date of installation.

The purchasing committee has the right to officially contact any of these hospitals and disqualify any offer where the feedback is negative in operation, after sales service or local agent performance.

- The vendor is responsible to ensure through official  documents that classified medical devices are manufactured in conformity with applicable quality system standards (ISO, IEC); (the international quality systems standard for medical device manufacturing ISO 13485) or ISO9001.

- With each offer, bidders must provide a formally endorsed document issued by the manufacturer stating that the bidder is the sole certified agent for the offered item.

- In all of the above cases (except 2.12) certificates must be formally endorsed by JFDA.

- Any vendor not submitting all required certificates will be eliminated.

- Official website of the manufacturer must be provided. The website must clearly demonstrate the history profile and manufacturing features of the company including the same offered model and its brochure.

In case of inconsistency between the information on the website with the provided information of the local agent regarding any offered product, the committee has the right to eliminate that product from the awarding process.

- Offered items should be from reputable well known manufacturers and excellent experience in the field and shall have multiple installations of the same offered model and/ or previous models in RMS main hospitals with at least two years of operation and excellent experience in operation, after sales service & local agent performance; otherwise the purchasing committee has the right to request any of the following:

- An evaluation certificate as mentioned in term 2.10.b.

- A sample of the same offered model at any time during the purchasing process, the sample unit shall be delivered within three weeks from the date of a written notification; any vendor rejects to deliver a sample unit will be eliminated. Evaluation duration will be determined by the purchasing committee.

Any offered item fail in the evaluation/assessment process will be rejected

- Vendors must specify the origin of the offered items and accessories in the technical offer.

- Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first half of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period shall be rejected.

- Goods are to be brand new; manufacturing date of the delivered products should not exceed 18 months from the date of the final award.

- Delivery period should be mentioned clearly; the purchasing committee has the right to reject any offer with delivery time exceeding 4 months.

- Warranty:

- Offers must include a full warranty for a period of a minimum of 24 months from the date of installation or 30 months from the date of receiving the items at the agreed location mentioned in the final order (whichever comes first).

Warranty must include corrective & preventive maintenance activities as per manufacturer recommendations including:

- Required spare parts (free of charge)

- Labour

- Hardware

- Software

- Rechargeable batteries

At the end of the warranty period, The supplier commits to implement final inspection of the submitted goods and submit the reports signed by the site chief engineer stating that the equipment are working properly as well as all preventive maintenance reports during the warranty period.

Warranty & installation are excluded for items mentioned in Attachment no.1 the purchasing committee has the right to exclude any other item not listed in the attachment and does not required installation and warranty.

- In the case where a delay in installation has occurred as a result of the supplier's dereliction and has exceeded a period of one month from the date of receiving the items, the approved warranty period will automatically be considered as 24 months from the date of installation.

- If at any time during the warranty period the item becomes inoperative due to a technical fault the item must then be repaired by the supplier /local agent within a period of 5 working days from written notification, warranty will be extended according to downtime period.

- If the delay exceeds 30 days the supplier must replace the item with a new identical functioning one (within the same delivery period mentioned in the final offer). In case the item was replaced by a new one, the warranty period mentioned in 9.a) above will start from the installation and commissioning date of the new item.

- Local agent/ supplier is committed to transport and install any awarded item during the warranty period on free of charge basis to any location inside the country whenever required by Royal Medical Services, this should not include pre-installation requirements in the new site.

- All offered items (main unit) should be fully designed, manufactured, and labelled by their real original manufacturer in which all related testing, research, development and approvals went through.

Any relabelled products for the main unit (white-label manufacturing, OEM, or repackaging) are rejected.

- One set of operation manual(s) and one set of service manual(s) including schematics and a spare-part list must be delivered with each unit, CD/DVD is acceptable. For large tenders, a certain agreed percentage of manuals per item may be agreed upon.

- Where applicable, pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include removal of old system(s), any civil work, electrical work or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new system(s).

- Power requirements: where applicable either single phase 220V, 50Hz or 3-phase 380V. Systems with external transformers are considered conforming only if clearly stated in the technical specifications.

- Technical offers must include clear original technical brochures/catalogues for all offered items.

- Offers must include fully detailed technical offers and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s), full technical description/specifications and any accessories or options included in the offer.

Offer must clearly indicate the origin of the offered item and the country where the manufacturer of the offered items is based.

- Compliance sheets must be as per the tabular format of the technical specifications in the tender documents, listing the required specifications on one column and a  Yes  or NO  response to each point in the adjacent column, with reference to page and line numbers in the relevant technical brochure. Offers not complying with this term shall be rejected.

- Qualifications and after sales service: The technical bid should contain all the necessary documents (training certificates) to prove the capability of the bidders for supplying and installing a trouble free equipment meeting the quality standards and technical specifications of the manufacturer and the ability of the bidders for providing efficient after sales service conducted by authorized certified biomedical service engineers with minimum 2 years experience in the same field.

- Accessories and consumables:

  - Any accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately.

All offered accessories and consumables must be approved by the manufacturer.

  - Technical offers must include a priced list for all accessories and consumables related to the required equipment as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) with prices fixed for a period of five years from the date of installation and commissioning with a maximum annual increase of 2%, any essential item not listed will be considered free of charge.

  - Accessories and consumables must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Where applicable, a start-up kit of accessories and consumable items must be provided with each system on a free-of-charge basis.

- Spare Parts:

  - Technical offers must include a comprehensive and priced list of all spare parts related to the awarded equipment (including rechargeable batteries) as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) valid for a minimum period of five years with a maximum annual increase of 2%, commencing at the end date of the warranty period, any essential item not listed will be considered free of charge.

  - Spare parts must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Prices of spare parts should be reasonable and will be taken into consideration during the purchasing process; the purchasing committee has the right to eliminate any offered item with unreasonable high prices of spare parts.

  - Delivery period of required spare parts should not exceed 2 months from the date of the final order.

- Spare parts, consumables and accessories availability must be guaranteed for a minimum period of ten years starting from the date of installation and commissioning.

- Local agent is committed to sign a service contract with Royal Medical Services to perform all relevant maintenance for the awarded equipment (whenever required by DRMS within a period that does not exceed one year from the end date of the warranty period and if mentioned clearly in the specifications) without exceeding the original mentioned cost and as per mentioned terms within minimum required period.

- Tender Awards:

  - For the final list of offers having a chance of winning the award, the awarding process shall be based on the accumulative value of both the offered item and its’ running cost (Total Cost of Ownership) over a period of seven years from the date of installation and commissioning. Only offers with the lowest total cost of ownership over a period of seven years from the date of installation and commissioning shall qualify for the award.

  - Running cost includes the value of consumables, accessories needed to operate the system over the same period as well as the cost of any service contract (where applicable).

- For PC/Laptop based systems:

  - Complete restoration medium (CD/DVD/etc.) of the operating system and the application software must be supplied.

  - Purchased licenses, software keys & dongles must be provided at any time during and after the warranty period on free of charge basis whenever required by DRMS for a minimum period of ten years starting from the date of installation and commissioning.

  - Where locally supplied computers, laptops & printers are offered, the offered model should be from well known manufacturer.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores, installation, pre-installation (if needed), training, commissioning, warranty and bringing the equipment into service.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- a. DRMS has the right to increase the awarded quantities by a percentage not exceeding 35% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

b. DRMS has the right to decrease the awarded quantities by a percentage not exceeding 50% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- a. Bidder is not allowed to submit more than one offer for the same item whether that’s solely or in coalition or partnership with other bidders.

b. Bidder is allowed to include within their offer optional alternatives for the same offered item from the same manufacturer.

- The supplier must furnish DRMS with a guarantee stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelve months from the date of final acceptance of the equipment by DRMS.

- Training:

  - For items where abroad service training courses are required in technical specifications, offers must include a certified service training program for at least 3 working days at a reputable center abroad recognized by the manufacturer for at least one biomedical engineer or biomedical technician; all costs inclusive, air tickets, , boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - For items where abroad user  training courses are required in technical specifications, offers must include a certified operator training program at a reputable center abroad recognized by the manufacturer for at least one operator; all costs inclusive, air tickets, boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - The period of the training courses must be according to the manufacturer’s program excluding traveling days and must be stated clearly in the technical offer.

  - Training Programs must conform to the following standards:

- User training must comprise understanding and use of operation manual(s), correct and safe operation of the equipment, as well as user preventive maintenance and calibration.

- Service training must comprise: theory, understanding and use of service manual(s), calibration, preventive maintenance procedure, and practical troubleshooting and repair exercises, and must be conducted by professional instructors employed or authorized by the system manufacturer.

- Service training must be conducted on a system of identical make, model, and configuration to that purchased by DRMS, and designated by the manufacturer or the local agent for training purposes.

- Certificates must be endorsed and officially sealed by the system manufacturer, legally empowering trainees to engage in user and service activities according to operation and service manual(s).

- Where applicable, offers must include an on-site user and service training.

30. For offers submitted in Jordanian dinars, Payment will be either by wire transfer or by cheque after final acceptance of goods. Any other way of payment will be rejected.

Attachment 1

| # | Equipment |
| --- | --- |
| 1 | Adjustable Auto Pipettes |
| 2 | Resuscitation Bag |
| 3 | Laryngoscope Set |
| 4 | Oxygen Flow meter wall type/ single |
| 5 | Regulator Suction with canister, wall vacuum outlet |
| 6 | Oxygen Regulator for Oxygen Cylinders |
| 7 | Pulse Oximeter, Finger type |
| 8 | Oxygen Cylinder |
| 9 | Doppler, portable |
| 10 | Diagnostic set, Portable |
| 11 | Direct Ophthalmoscope, Portable |
| 12 | Otoscope, Portable |
| 13 | Air Mattress System, homecare |
| 14 | Stethoscope |
| 15 | Aneroid Sphygmomanometer |
| 16 | Video Assisted Laryngoscope, portable |
| 17 | Scale, Manual |
| 18 | Wood's Light |
| 19 | Cough Pressure, Normal Saline |
| 20 | Rehabilitation Walking Parallel Bars, non-powered |
| 21 | Therapy Mat |
| 22 | Medical Ball All Size |
| 23 | Dumbbells Rack with complete set of dumbbells |
| 24 | Crutches |
| 25 | Shoulder wheel |
| 26 | Mobile Mirror |
| 27 | Cuff Weights |
| 28 | Walker, different sizes |
| 29 | Patient Elbow Stick |
| 30 | Rehabilitation Training Ladder |
| 31 | Rehabilitation Suspension Frame |
| 32 | Exercise Band All Size (Theraband) |
| 33 | Lens trial set |
| 34 | Wheel Chair |
| 35 | Commode Chair |
| 36 | Bassinet (Baby Cot) |
| 37 | Resuscitation Cart (Crash Cart) |
| 38 | Medication Cart |
| 39 | Cart, Drawers |
| 40 | Examination Couch, Manual |
| 41 | Gynaecology Examination Table, Manual |
| 42 | Examination Table, Neonates, Manual |
| 43 | Intravenous Pole, Mobile Stand |
| 44 | DDA Cabinet |
| 45 | Stainless steel Multipurpose Trolley |
| 46 | Cabinet, Instrument, Operation Theatres |
| 47 | Dressing Cart |
| 48 | Stainless steel wire shelving unit |
| 49 | Paper Trolley |
| 50 | Stainless Steel Sink (Clean up counter) |
| 51 | Scopes Cabinet |
| 52 | Mayo Table |
| 53 | Table, Instrument |
| 54 | Stool, Adjustable, Doctor |
| 55 | Stool, Adjustable, Operation Theatres |
| 56 | Carts, linen/laundry, soiled, Double |
| 57 | Closed distribution trolley |
| 58 | Step Ladder, Conductive, Double |
| 59 | Step, Surgeon, Single |
| 60 | Kick Bucket |
| 61 | Mobile Stand for Oxygen Cylinder |
| 62 | Stainless Steel Wire Basket, 1 STU |
| 63 | Cart, Plaster |

Annex 5: General Surgery & Transplant & Gynecology & Pain Management

A: General Surgery

| Item A-1 | Laparoscopic Surgical Tower | Qty. (8) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| A | Durable, heavy duty construction, reliable with superior image quality |  |  |
| B | Camera system (4k Resolution): Provides electronic shutter function, automatic brightness control and optical zoom |  |  |
|  | - Resolution (4k): a minimum of 3840*2160. |  |  |
|  | - Compact, lightweight, easy to grip camera head with low heat generation |  |  |
|  | - Focal Distance with zoom lens, 16-28 mm or better |  |  |
|  | - Gas sterilizable, autoclavable is preferable |  |  |
|  | - Immersible in disinfections liquid. |  |  |
|  | - Universal coupling mechanism to fit all standard endoscopes. |  |  |
|  | - Camera cable 3 m or more |  |  |
|  | - Camera controller (4k resolution) |  |  |
|  | - Picture in picture mode |  |  |
| C | Light Source (either standalone device or integrated on camera control unit): |  |  |
|  | - LED Light source (light intensity should be equivalent to 300 Watt xenon lamp) |  |  |
|  | - Working hour ≥ 30,000 hours. |  |  |
|  | - Automatic and manual adjustment of light intensity. |  |  |
|  | - Front panel hygienic with convenient indicators for ease of operation. |  |  |
|  | - Flexible fibre optic cable preferably autoclavable and not less than 230 mm. |  |  |
| D | CO2 Insufflator: |  |  |
|  | - Sealed hygienic front touch panel. |  |  |
|  | - Maximum flow rate From 40- 50 L/MIN. |  |  |
|  | - Display of gas consumption, flow rate and intra-abdominal pressure |  |  |
|  | - Gas heating to body temperature (at 37 C). |  |  |
|  | - Audible and visual alarms. |  |  |
|  | - Filter on the CO2 output to prevent cross-contamination. |  |  |
|  | - Smoke evacuator that is either integrated or quoted as a separate unit. |  |  |
|  | - Accessories: Supplied with all necessary accessories including tubing set, CO2, hoses, connectors and filter. |  |  |
| E | Suction/ Irrigation combined unit: |  |  |
|  | - Complete with all standard sets of tubes for the unit. |  |  |
|  | - Roller pump for irrigation (Pressure: adjustable up to 200 mmHg or better). |  |  |
|  | - Flow Rate: adjustable up to 1 litre//min or better. |  |  |
|  | - Adjustable vacuum pressure. |  |  |
|  | - Complete with all standard sets of tubes for the unit. |  |  |
| F | LCD / LED Colour Monitor: Qty. (1) |  |  |
|  | - Medical grade type that is supplied by the same manufacturer |  |  |
|  | - At least of 32 inch size that is mounted on the mobile cart |  |  |
|  | - Resolution: 4k resolution minimum of 3840*2160 |  |  |
|  | - DVI compatible. |  |  |
| G | Ability to record images and videos on Medical grade DVD recorder or USB flash memory |  |  |
| H | Telescopes, 4 K Resolution, Qty. 1 each with each tower: |  |  |
|  | - TELESCOPE 30° (4K) Diameter 10 mm |  |  |
|  | - TELESCOPE 0° (4K) Diameter 10 mm |  |  |
|  | - TELESCOPE 30° (4K) Diameter 5 mm |  |  |
|  | - TELESCOPE 0° (4K) Diameter 5 mm |  |  |
| I | ICG option should be available and quoted separately, 4K Camera, Camera controller, Light source & telescopes should be ICG compatible |  |  |
| J | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for the subcomponents for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |
| K | Abroad service training for 2 biomedical engineers/ technicians as per special terms |  |  |

| Item A-2 | Intra-Operative Nerve Locator/Stimulator, Head & Neck | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | From a reputable well known manufacturer; The committee has the right to request for a sample of the same offered model at any time during the purchasing process, the sample unit shall be delivered within three weeks from the date of a written notification; any vendor rejects to deliver a sample unit will be eliminated. |  |  |
| 2 | The system must be the latest model that is Suitable for endocrine surgeries for monitoring of Recurrent Laryngeal Nerve |  |  |
| 3 | 4- Channels EMG monitoring capability |  |  |
| 4 | # Audible and visual real-time feedback upon nerve irritation to aid in intraoperative decision making |  |  |
| 5 | # Large coloured touch screen |  |  |
| 6 | # Patient Interface (wired or wireless) between electrodes and the screen |  |  |
| 7 | # Stimulator Probe, monopolar |  |  |
| 8 | # Muting detector probe to reject noise generated from electrosurgical units |  |  |
| 9 | # EMG recording capability |  |  |
| 10 | # The unit should be small in size and fit easily on the surgical pendant shelf |  |  |
| 11 | # Printer should be quoted separately |  |  |
| 12 | # All consumables including (probes and sensors) should be listed and quoted separately |  |  |
| 13 | # Options to be quoted separately |  |  |
| 14 | # All other accessories to be priced separately |  |  |

B: Transplant

| Item B-1 | Ultrasonic Surgical Aspirator | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Cavity Aspiration Technology for the fragmentation and removal of soft tissues, Bidder must submit a system (with the same offered model) to be assessed and /or evaluated by RMS when required by the purchasing committee during one month from a written notification, offers which fail the evaluation/assessment process will be excluded from the tender |  |  |
| 2 | Control Unit: |  |  |
|  | - Microprocessor controlled. |  |  |
|  | - Controls for irrigation rate, vacuum limits and ultrasound power. |  |  |
|  | - Display for the status of irrigation, vacuum and ultrasound. |  |  |
|  | - Ultrasonic frequency 22-35 KHz |  |  |
|  | - Integrated or external aspiration |  |  |
|  | - Integrated irrigation unit. |  |  |
|  | - Digital display of settings. |  |  |
|  | - Noiseless operation. |  |  |
|  | - Display of errors messages |  |  |
| 3 | Hand-pieces: |  |  |
|  | - hand-piece22-35 kHz all options should be quoted separately Qty.(2) |  |  |
|  | - Light weight |  |  |
|  | - Easily connected to the control unit |  |  |
|  | - Autoclavable |  |  |
|  | - No cooling required. |  |  |
|  | - All Standard accessories should be included |  |  |
| 4 | Footswitch Qty. (1) |  |  |
| 5 | Tips (with sleeves), all sizes should be priced separately Qty.(60) |  |  |
| 6 | Sterilization Tray Qty. (2) |  |  |
| 7 | Tubing set (suction& irrigation) & suction filter Qty. (30) |  |  |
| 8 | Reusable suction canister Qty. (1) |  |  |
| 9 | Original Trolley Qty. (1) |  |  |
| 10 | IV Pole Qty. (1) |  |  |

| Item B-2 | Rapid Infusion System | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Microprocessor controlled with high resolution large and easy to read LCD display and hygienic flat touch control panel |  |  |
| 2 | Shall be used for rapid warming and infusing of blood/fluid in case of trauma, emergency, sever bleeding etc. |  |  |
| 3 | Protection against fluid spillage (IPX2 or better) |  |  |
| 4 | With internal rechargeable battery |  |  |
| 5 | With user selectable infusion modes |  |  |
| 6 | Shall include mounting clamp to standard IV poles |  |  |
| 7 | Extremely high flow rate infusion: |  |  |
|  | - Adjustable ≈ 20 - 1000 ml/minute |  |  |
|  | - Tolerance ≈ ± 10% from 20 – 1000 ml/min |  |  |
|  | - With integrated blood/fluid warmer (up to 40 C° approximately) |  |  |
|  | - Adjustable and programmable bolus delivery |  |  |
| 8 | Safety features: |  |  |
|  | - Aluminium free fluid path |  |  |
|  | - Occlusion pressure monitoring |  |  |
|  | - Ultrasonic air bubble detector (with alarm and pump shutdown) |  |  |
|  | - Audio/visual alarming system |  |  |
| 9 | IV administration sets: Disposable latex free IV set price should be offered separately with fixed price as per special terms, the price will be part of financial evaluation of the machine |  |  |
| 10 | Accessories: Complete with basic accessory and any optional accessories should be priced separately |  |  |

| Item B-3 | High-End Ultrasound Machine Dedicated for Intraoperative and Laparoscopic Use | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | High-end Ultrasound Machine dedicated for intraoperative and laparoscopic applications. The system must meet or transcend the minimum requirements listed below. |  |  |
| A | # Application: Abdominal, intraoperative, laparoscopic and hepatic surgery |  |  |
| B | ## Configuration: |  |  |
|  | - Mobile cart/built-in trolley type. |  |  |
|  | - Four full swivel castors with brakes |  |  |
|  | - Full control keyboard |  |  |
|  | - High-resolution LCD: ≥ 21” |  |  |
|  | - Storage for transducer cables |  |  |
|  | - Auxiliary storage compartments/drawers |  |  |
|  | - Minimum of 4 active transducer connectors |  |  |
|  | - User-adjustable presets |  |  |
|  | - Adjustable control panel: up/down & rotate |  |  |
|  | - Adjustable monitor position: up/down, rotate, swivel, & tilt |  |  |
| C | General Features: |  |  |
|  | - Gray-scale levels: 256 |  |  |
|  | - Power up time: ≤ 2 minutes |  |  |
|  | - Self-diagnostic routine during power up |  |  |
|  | - Touch screen capability |  |  |
|  | - Acquisition frame rate: ≥ 1600 fps |  |  |
|  | - Simultaneous processing bandwidth capability: 1-18MHz |  |  |
|  | - No. of digitally processed channels: ≥ 400000 |  |  |
|  | - Input dynamic range: : ≥ 240Db |  |  |
|  | - Cine loop: ≥ 1000 frames |  |  |
|  | - Automatic real-time Doppler trace capability including calculation and display of user-selected measurements |  |  |
|  | - Real-time Pan/Zoom |  |  |
|  | - Frozen image Pan/Zoom |  |  |
|  | - Automatic optimization of base line and scale |  |  |
|  | - Local HDD patient search through name and/or ID |  |  |
|  | - Full DICOM 3.0, RIS, HIS, PACS Support |  |  |
|  | - Raw data processing capability Or post processing or equivalent |  |  |
| D | Image Storage: |  |  |
|  | - Capacity of stored images HDD: ≥ at least 300 GB |  |  |
|  | - Removable Storage: CD/DVD, USB |  |  |
| E | Multi-frequency Transducers: |  |  |
|  | - Linear Array:6 - 18 MHz with biopsy guides |  |  |
|  | - Curved Array:2 -5 MHz with biopsy guides |  |  |
|  | - Laparoscopic Probe: 5 - 10 MHz, articulated tip |  |  |
|  | - Linear Array:3 - 8 MHz for vascular application, with biopsy guides |  |  |
|  | - Intraoperative Probe: Wide Band, 5-14 MHz, to be Sterilizable |  |  |
|  | - Intraoperative Robotic Probe: Option to be available and priced separately |  |  |
| F | Imaging Modes: |  |  |
|  | - 2D |  |  |
|  | - B-mode |  |  |
|  | - Tissue harmonic imaging |  |  |
|  | - M-mode and colour M-mode |  |  |
|  | - Dual imaging |  |  |
|  | - Doppler imaging |  |  |
|  | - Color Doppler imaging |  |  |
|  | - Color Power Angio Imaging |  |  |
|  | - Pulsed Wave (PW) Doppler |  |  |
|  | - Simultaneous PW Doppler and 2D (Dual Mode) |  |  |
|  | - Simultaneous 2D, color Doppler, and PW Doppler (Triple Mode) |  |  |
|  | - Contrast imaging |  |  |
|  | - Compounding imaging |  |  |
|  | - Speckle reduction (Noise reduction imaging) |  |  |
| G | All available standard & optional features, packages, & accessories must be listed and priced separately |  |  |

C: Gynecology

| Item C-1 | Gynaecology & Hysteroscopy Surgical Laparoscopic Tower | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| A | Durable, heavy duty construction, reliable with superior image quality |  |  |
| B | Camera system (4k Resolution) : Provides electronic shutter function, automatic brightness control and optical zoom |  |  |
|  | - Resolution (4k): a minimum of 3840*2160. |  |  |
|  | - Compact, lightweight, easy to grip with low heat generation |  |  |
|  | - Focal Distance with zoom lens, 16-28 mm or better |  |  |
|  | - Gas sterilizable, autoclavable is preferable |  |  |
|  | - Immersible in disinfections liquid. |  |  |
|  | - Universal coupling mechanism to fit all standard endoscopes. |  |  |
|  | - Camera cable 3 m or more |  |  |
|  | - Camera controller (4k resolution) |  |  |
|  | - Picture in picture mode |  |  |
| C | Light Source (either standalone device or integrated on camera control unit): |  |  |
|  | - LED Light source (light intensity should be equivalent to 300 Watt xenon lamp) |  |  |
|  | - Working hour ≥ 30,000 hours. |  |  |
|  | - Automatic and manual adjustment of light intensity. |  |  |
|  | - Front panel hygienic with convenient indicators for ease of operation. |  |  |
|  | - Flexible fibre optic cable preferably autoclavable and not less than 230 mm. |  |  |
| D | CO2 Insufflator: |  |  |
|  | - Sealed hygienic front touch panel. |  |  |
|  | - Display of settings. |  |  |
|  | - Maximum flow rate From 40- 50 L/MIN. |  |  |
|  | - Display of gas consumption, flow rate and intra-abdominal pressure |  |  |
|  | - Gas heating to body temperature (at 37 C). |  |  |
|  | - Audible and visual alarms. |  |  |
|  | - Filter on the CO2 output to prevent cross-contamination. |  |  |
|  | - Smoke evacuator that is either integrated or quoted as a separate unit. |  |  |
| E | Suction/ Irrigation combined unit: |  |  |
|  | - Complete with all standard sets of tubes for the unit. |  |  |
|  | - Roller pump for irrigation (Pressure: adjustable up to 150 mmHg or better). |  |  |
|  | - Flow Rate: adjustable up to 0.5 litre//min or better. |  |  |
|  | - Adjustable vacuum pressure. |  |  |
|  | - Complete hysteroscopy set for irrigation system |  |  |
| F | LCD / LED Colour Monitor: Qty. (1) |  |  |
|  | - Medical grade type that is supplied by the same manufacturer |  |  |
|  | - At least of 32 inch size that is mounted on the mobile cart |  |  |
|  | - Resolution: 4k resolution minimum of 3840*2160 |  |  |
|  | - DVI compatible. |  |  |
| G | Ability to record images and videos on Medical grade DVD recorder or USB flash memory |  |  |
| H | Telescopes, 4 K Resolution, Qty. 1 each with each tower: |  |  |
|  | - Straight forward telescope 0º, 4K resolution, enlarged view, diameter 10 mm, length 30-31 cm, autoclavable |  |  |
|  | - Forward oblique telescope 30º, 4K resolution, enlarged view, diameter 10 mm, length 30-31 cm |  |  |
|  | - Straight forward telescope 0º, 4K resolution, enlarged view, diameter 5 mm, length 29-30 cm, autoclavable |  |  |
|  | - Forward oblique telescope 30º, 4K resolution, enlarged view, diameter 5 mm, length 29-30 cm |  |  |
| I | ICG option should be available and quoted separately, 4K Camera, Camera controller, Light source & telescopes should be ICG compatible |  |  |
| J | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for the subcomponents for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |

D: Thoracic

| Item D-1 | Portable Video Mediastinoscope System | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Durable, heavy duty construction, reliable with superior image quality |  |  |
| 2 | Portable imaging system (console): |  |  |
|  | - Must be portable and compact (one unit) |  |  |
|  | - Full HD resolution touch screen monitor with size of at least 18 inch |  |  |
|  | - Built in LED light source |  |  |
|  | - Automatic and manual light control |  |  |
|  | - Built in documentation system |  |  |
|  | - Can be connected to another monitor |  |  |
|  | - Picture in picture mode |  |  |
|  | - Hygienic |  |  |
|  | - Internal storage 50 GB or more |  |  |
|  | - USB storage of HD images and videos |  |  |
|  | - Must be Light weight 10 KG or less |  |  |
| 3 | Video mediastinoscope blade |  |  |
|  | - Compact Video mediastinoscope for imaging system (item 1) |  |  |
|  | - Integrated CMOS Chip on the tip FHD Camera |  |  |
|  | - Build in LED light source |  |  |
|  | - Reusable and autoclavable type |  |  |
|  | - Minimum of 3 buttons to control the system |  |  |
|  | - Blade with continuous lateral slit |  |  |
|  | - Blade Length 15 cm approximately |  |  |
|  | - Made from stainless-steel or titanium |  |  |
|  | - Cable length 290 cm or more |  |  |
| 4 | Wire Tray for cleaning and sterilization |  |  |
| 5 | Silicon grip insert. |  |  |
| 6 | All standard accessories should be included |  |  |
| 7 | Option: mobile trolley with should be priced separately |  |  |

| Item D-2 | Electrosurgical Unit & Argon Plasma for Thoracic Surgery | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Name of Manufacturer |  |  |  |  |  |
| Model/ catalogue number |  |  |  |  |  |
| Country of Origin for the offered model |  |  |  |  |  |
| Country where the manufacturer is based |  |  |  |  |  |
| Delivery time |  |  |  |  |  |
| Full warranty period |  |  |  |  |  |
|  | Minimum Requirements |  | Compliance )Y/N), Notes | Brochure Page No. |
| A | Electrosurgical unit: |  |  |  |
|  | - Monopolar cut mode with at least 200 watt power |  |  |  |
|  | - Monopolar coagulation mode with at least 120 watt power |  |  |  |
|  | - Bipolar coagulation with at least 90 watt power |  |  |  |
|  | - Automatic output dosage control |  |  |  |
|  | - Monopolar, bipolar , neutral electrode plugs must be compatible with more than one cable type and manufacturer |  |  |  |
|  | - Special cutting mode for polypectomy |  |  |  |
|  | - Special cutting mode for needle instrument |  |  |  |
|  | - LCD display screen |  |  |  |
|  | - Selectable effects for cut and coagulation power |  |  |  |
|  | - Output frequency ≥ 350 KHz |  |  |  |
|  | - Leakage current less than 20 mA |  |  |  |
|  | - Accessories: Monopolar adjustable adapter, patient plate cable, disposable patient plate |  |  |  |
| B | Argon Plasma: |  |  |  |
|  | - Non–contact procedure |  |  |  |
|  | - Gas flow: 0.1-8 L approximately |  |  |  |
|  | - Flashing flow: 3 sec approximately |  |  |  |
|  | - Flexible Argon Probe 1.5 mm diameter, Length 1.5 meter for flexible bronchoscopy with integrated filter, 10 Pcs |  |  |  |
|  | - APC (Argon Plasma coagulation) Rigid applicator with integrated filter, 5 Pcs |  |  |  |
| C | Cryo-Therapy Unit for pulmonary applications to be quoted separately |  |  |  |
| D | Original trolley should be quoted separately |  |  |  |
| E | Twin footswitch |  |  |  |
| F | Accessories to be quoted separately |  |  |  |

| Item D-3 | Therapeutic Fibre-Optic Bronchoscope | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Name of Manufacturer |  |  |  |  |  |
| Model/ catalogue number |  |  |  |  |  |
| Country of Origin for the offered model |  |  |  |  |  |
| Country where the manufacturer is based |  |  |  |  |  |
| Delivery time |  |  |  |  |  |
| Full warranty period |  |  |  |  |  |
|  | Minimum Requirements |  | Compliance )Y/N), Notes | Brochure Page No. |
| A | Therapeutic fiber-optic flexible bronchoscope of the following specifications: |  |  |  |
|  | - The entire bronchoscope must be submersible. |  |  |  |
|  | - Optical system: |  |  |  |
|  | - Viewing direction: ( forward viewing 0 º) |  |  |  |
|  | - Field of view ≥ 110º |  |  |  |
|  | - Observation range: (3 - 50) mm or better |  |  |  |
|  | - Insertion tube section: |  |  |  |
|  | - insertion tube diameter (5.7-6.6)mm |  |  |  |
|  | - working length ≥ 540 mm |  |  |  |
|  | - One working channel with a diameter: (2.6-3) mm |  |  |  |
|  | - Tip deflection: |  |  |  |
|  | - Up : (180 º) |  |  |  |
|  | - Down ≥ 100 º |  |  |  |
|  | - Supports coupling with an external light source through a fiber optic cable; required coupling adapters shall be provided |  |  |  |
|  | - Equipped with a universal coupling, enabling compatibility with a wide range of external camera heads. |  |  |  |
|  | - Compatible with ETO & liquid reprocessing methods. |  |  |  |
|  | - Case with complete accessories kit should be listed separately. |  |  |  |
| B | Leakage tester Qty. (1) |  |  |  |

| Item D-4 | Sternum Saw, Battery Type | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Name of Manufacturer |  |  |  |  |  |
| Model/ catalogue number |  |  |  |  |  |
| Country of Origin for the offered model |  |  |  |  |  |
| Country where the manufacturer is based |  |  |  |  |  |
| Delivery time |  |  |  |  |  |
| Full warranty period |  |  |  |  |  |
|  | Minimum Requirements |  | Compliance )Y/N), Notes | Brochure Page No. |
| 1 | Durable, heavy duty construction, high performance & reliable |  |  |  |
| 2 | Pistol grip hand piece bone cutter dedicated to perform median sternotomy |  |  |  |
| 3 | Ergonomic, light weight, small in size and easy to handle |  |  |  |
| 4 | Configuration: battery type with complete accessories, blades, guards, battery packs & battery charger |  |  |  |
| 5 | Construction: stainless steel or aluminium |  |  |  |
| 6 | Trigger switch which controls the variable speed operation of the handpiece according to the pressure applied to the trigger switch |  |  |  |
| 7 | It should have reciprocating blade motion |  |  |  |
| 8 | It should have vibration control with no or minimal vibration |  |  |  |
| 9 | It should have high precision and quick cutting |  |  |  |
| 10 | Reliable, quick and easy replacements of blades (key or key-less chuck mounting mechanism) |  |  |  |
| 11 | It should have facility for intra operative replacement of battery pack |  |  |  |
| 12 | It should have safe mode |  |  |  |
| 13 | CPM ≥ 13000 |  |  |  |
| 14 | Simple daily maintenance. |  |  |  |
| 15 | All accessories and consumables including sternum guard needed to perform median sternotomy should be listed and priced separately |  |  |  |
| 16 | Batteries: |  |  |  |
|  | - Battery continuous operation: at least 20 min |  |  |  |
|  | - Two batteries for each hand piece with autoclavable aseptic housing and aseptic shield, all should be priced separately, price should be fixed as per special terms |  |  |  |
|  | - One battery charger to be included with at least 4 battery slots with independent battery charge status display |  |  |  |
| 17 | Sterilization container for the whole system including accessories |  |  |  |
| 18 | Handpiece, housing, shield and sternum guard must be easily sterilizable by autoclaving (at standard 134°C cycle with standard sterilization, pre-vacuum and drying periods) and plasma sterilization |  |  |  |
| 19 | In case all offers are not compliant with special term (19) considering the spare parts; preference will be given to the offer submitting the best alternative for the term after the warranty period (i.e. fixed exchange price or fixed maintenance cost for a period of at least 5 years after the warranty period with a maximum annual increase of 2% annual increase) |  |  |  |

| Item D-5 | Surgical Laser System, ND-YAG | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | It should be stated clearly in the original catalogue that the offered model is optimized for wide range of thoracic surgical applications in both open and minimally invasive surgeries with precise resection, coagulation and tissue sealing capabilities and excellent haemostasis. |  |  |
| 2 | # Treatment Laser type: Nd:YAG (Neodymium-doped Yttrium Aluminum Garnet) |  |  |
| 3 | # Aiming Laser Type: Visible diode laser |  |  |
| 4 | # Wavelength ≈ 1320 nm ± 10 nm |  |  |
| 5 | # Energy Output: Adjustable from 1 to at least 100 watts |  |  |
| 6 | # Treatment Modes of operation: - Continuous Wave (CW) - Pulsed Wave (PW) |  |  |
| 7 | # Fiber Optic Delivery System: - Compatible with thoracoscope and bronchoscope - Reusable and autoclavable flexible fiber, different diameters should be quoted separately |  |  |
| 8 | # Built in smoke evacuator |  |  |
| 9 | # Efficient built in air cooling system |  |  |
| 10 | # Large LCD digital touch screen display ≥ 7 inch with clear indicators for laser status |  |  |
| 11 | # Noise level should not exceed 60 dB(A) |  |  |
| 12 | # Large emergency stop to stop laser emission immediately |  |  |
| 13 | # Mobile type on large castors |  |  |
| 14 | # Footswitch control: Dual pedal |  |  |
| 15 | # All accessories and consumables (focusing hand-piece, fibers, fiber preparation set, fiber holder, tips, adapters, filters, funnels, suction tubes, air hoses, laser safety goggles and sterilization tray) and any other required accessories or consumables needed for the above treatments should be quoted separately. |  |  |

E: Pain Management

| Item E-1 | Cryo-Analgesia System for Pain Treatment | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Name of Manufacturer |  |  |  |  |  |
| Model/ catalogue number |  |  |  |  |  |
| Country of Origin for the offered model |  |  |  |  |  |
| Country where the manufacturer is based |  |  |  |  |  |
| Delivery time |  |  |  |  |  |
| Full warranty period |  |  |  |  |  |
|  | Minimum Requirements |  | Compliance )Y/N), Notes | Brochure Page No. |
| 1 | Application: for chronic pain relieving (Low back pain, upper and lower extremities, craniofacial, chest wall, abdominal and pelvic, sacroiliac joint, knee, etc…) by inserting a small probe to freeze the target nerve to temporarily block nerve conduction and facilitating complete regeneration of the structure and function of the affected nerve. |  |  |  |
| 2 | To work on liquid nitrogen or carbon dioxide |  |  |  |
| 3 | Adjustable pressure and gas flow rate |  |  |  |
| 4 | Freezing temperature within the range 75-90 °C |  |  |  |
| 5 | Coloured touch screen of size not less than 10” with large numeric readouts of cryoprobe temperature, cylinder pressure, gas flow inside cryoprobe and procedure time |  |  |  |
| 6 | Foot switch |  |  |  |
| 7 | Simple, easy to use with and quick setup and user friendly interface |  |  |  |
| 8 | Customizable presets and settings for individual physician users and procedures |  |  |  |
| 9 | All accessories needed to run the equipment should be included |  |  |  |
| 10 | Bidders are required to offer all available consumables including cryoprobes and any required items needed for pain management (reusable & disposable), it should be listed and priced separately |  |  |  |

| Item E-2 | C-ARM | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | A state of the art digital Mobile C-Arm fluoroscopy system is required to provide excellent image quality and minimal dose for a wide range of applications. The system must comprise sufficiently sophisticated hardware and software platforms to be able to host a wide range of real-time and post processing image enhancement techniques, the system must meet or transcend the following specifications: |  |  |
| 1 | Application: Orthopaedics, Trauma surgery, General surgical, Pain Management, interventional procedures, Vascular surgery; Angiography and DSA |  |  |
| 2 | ## Physical Configuration: Two wheeled units with smooth manual steering system: one supporting the C-Arm and console and the other supporting monitors, image processing, recording devices etc. |  |  |
| 3 | ## System backbone: Microprocessor control over all subsystems with an elaborate self-test and error code scheme. |  |  |
| 4 | Control Console: hygienic touch type controls with easy to read, user friendly annotation. |  |  |
| 5 | X-Ray Tube: |  |  |
|  | - Rotating anode type. |  |  |
|  | - Anode heat Capacity, kHU ≥ 300 |  |  |
|  | - Anode cooling, kHU/min 50 |  |  |
|  | - Dual focus, size: ≤ 0.6 each |  |  |
| 6 | X-Ray Generator: |  |  |
|  | - Ripple-free high-frequency converter type. |  |  |
|  | - Power rating, kW@100 kVp: ≥ 12 kW. |  |  |
|  | - Radiographic Mode: - kV range: 40 to 120 kV or better in steps of 1 kV. - mAs range: up to 60 mA or better - AEC |  |  |
|  | - Fluoroscopic Mode: - kV range: 40 to 110 kV or better - mA range: up to 8 mA or better - Pulsed fluoroscopy: ~ 1 to 25 pulses per sec. - Digital Snapshot mode |  |  |
| 7 | Flat panel detector |  |  |
|  | - Size not less than (26 cm × 26 cm) |  |  |
|  | - Pixel size ≤ 200 µm or less |  |  |
|  | - CMOS flat panel detector type |  |  |
|  | - DQE ≥ 70% |  |  |
| 8 | TV Monitor: |  |  |
|  | - Dual high resolution displays of at least 18 inches coloured LCD monitors, or one 32 inches coloured LCD monitor |  |  |
|  | - Mounted on an articulating arm |  |  |
| 9 | Image Acquisition: |  |  |
|  | - Image matrix: ~ 1024x1024. |  |  |
|  | - Collimation: Dual Leaf and Iris type. |  |  |
|  | - Image reversal and rotation capability. |  |  |
|  | - Magnification (detector zoom) |  |  |
| 10 | Image Processing and Storage: |  |  |
|  | - Image storage matrix: ~ 1024x1024. |  |  |
|  | - Patient data registration system. |  |  |
|  | - Last image hold. |  |  |
|  | - Cine replay. |  |  |
|  | - Digital subtraction. |  |  |
|  | - CD/DVD R archiving. |  |  |
|  | - Full DICOM 3.0 connectivity. |  |  |
|  | - Kindly quote for all available software packages (e.g. DSA, Roadmap, Zoom, Image enhancement, dedicated vascular package, dedicated Ortho Package …. etc |  |  |
| 11 | C-Arm: |  |  |
|  | - Free Space: ≈ 70cm. |  |  |
|  | - Depth: ≈ 60cm. |  |  |
|  | - Vertical Travel: ≈ 40cm. |  |  |
|  | - Orbital and pivot Rotation, Horizontal and panning motion, and reverse position. |  |  |
| 12 | A list of standard accessories for the offered model (optional). |  |  |
| 13 | Offer must include the price for a comprehensive service contract (including all replacement parts and x-ray tubes) valid for a period of (8) years starting at the end of the 24 month warranty period |  |  |
| 14 | Abroad service training for 1 biomedical engineer/ technician as per special terms |  |  |

SPECIAL TERMS

- Offers not complying with any of the special terms or the technical specifications shall be considered non-conforming with tender requirements.

- Any vendor providing FORGED documents shall be disqualified from the current tender and banned from participating in any future RMS tenders.

- Wherever term “based” is mentioned it refers to the country where the manufacturing company is founded & established.

- All equipment must be the most recently released model/version which is equal to or higher than the range of the specifications of the required system (low, mid or high) and equal to or higher than the level of technology and required options mentioned in  the technical specifications.

- Required certificates (must be submitted with the technical offer):

- FDA clearance (510K) for equipment of USA based companies.

- MDR (EU) 2017/745 certificate for equipment of EU (European Union) based companies.

- UKCA certificate For Equipment of Great Britain based companies (England, Scotland, Northern Ireland, and Wales).

- ARTG (Australian Register of Therapeutic Goods) certification or approval for Australian and New Zealand based companies.

- Establishment License with the relevant Device License issued by the Therapeutic Products Directorate for Equipment of Canadian based companies.

- PMDA (Pharmaceuticals and Medical Devices Agency) certification or approval for equipment of Japanese based companies.

- Swissmedic (Swiss agency for therapeutic products) certification or approval for equipment of Swiss based companies.

- Norwegian Medicines Agency certification or approval for equipment of Norwegian based companies.

- Only for class Ι medical equipment manufactured by companies based in one of the countries mentioned above, submission of either one of the certificates mentioned above or a free sale certificate in any of these countries shall be accepted.

- For the following equipment:

- All X-ray equipment, MRI, ultrasound, and nuclear medicine systems (regardless where the manufacturing companies are based).

- Where the manufacturing companies are based in other origins than the mentioned in terms 2.1 - 2.8.

The following are required:

- At least two of the certificates mentioned above, one of which has to be FDA clearance (510K) (Only for class Ι medical equipment submission of certificates mentioned in 2.9 shall be accepted).

- Evaluation certificate from the Royal Medical Services for the same offered model with at least 80% passing grade.

If the evaluation is not applicable (based on purchasing committee perspective) bidder should submit a list of installation basis of the same offered model and/ or previous models in at least two of the following hospitals (King Hussein Cancer Center, National Center for Diabetes Endocrinology and Genetic Diseases, Jordan University Hospital or King Abdullah University Hospital) with at least three years of operation, list should include: Name of hospital, Model installed, Quantity, and date of installation.

The purchasing committee has the right to officially contact any of these hospitals and disqualify any offer where the feedback is negative in operation, after sales service or local agent performance.

- The vendor is responsible to ensure through official  documents that classified medical devices are manufactured in conformity with applicable quality system standards (ISO, IEC); (the international quality systems standard for medical device manufacturing ISO 13485) or ISO9001.

- With each offer, bidders must provide a formally endorsed document issued by the manufacturer stating that the bidder is the sole certified agent for the offered item.

- In all of the above cases (except 2.12) certificates must be formally endorsed by JFDA.

- Any vendor not submitting all required certificates will be eliminated.

- Official website of the manufacturer must be provided. The website must clearly demonstrate the history profile and manufacturing features of the company including the same offered model and its brochure.

In case of inconsistency between the information on the website with the provided information of the local agent regarding any offered product, the committee has the right to eliminate that product from the awarding process.

- Offered items should be from reputable well known manufacturers and excellent experience in the field and shall have multiple installations of the same offered model and/ or previous models in RMS main hospitals with at least two years of operation and excellent experience in operation, after sales service & local agent performance; otherwise the purchasing committee has the right to request any of the following:

- An evaluation certificate as mentioned in term 2.10.b.

- A sample of the same offered model at any time during the purchasing process, the sample unit shall be delivered within three weeks from the date of a written notification; any vendor rejects to deliver a sample unit will be eliminated. Evaluation duration will be determined by the purchasing committee.

Any offered item fail in the evaluation/assessment process will be rejected

- Vendors must specify the origin of the offered items and accessories in the technical offer.

- Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first half of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period shall be rejected.

- Goods are to be brand new; manufacturing date of the delivered products should not exceed 18 months from the date of the final award.

- Delivery period should be mentioned clearly; the purchasing committee has the right to reject any offer with delivery time exceeding 4 months.

- Warranty:

- Offers must include a full warranty for a period of a minimum of 24 months from the date of installation or 30 months from the date of receiving the items at the agreed location mentioned in the final order (whichever comes first).

Warranty must include corrective & preventive maintenance activities as per manufacturer recommendations including:

- Required spare parts (free of charge)

- Labour

- Hardware

- Software

- Rechargeable batteries

At the end of the warranty period, The supplier commits to implement final inspection of the submitted goods and submit the reports signed by the site chief engineer stating that the equipment are working properly as well as all preventive maintenance reports during the warranty period.

Warranty & installation are excluded for items mentioned in Attachment no.1 the purchasing committee has the right to exclude any other item not listed in the attachment and does not required installation and warranty.

- In the case where a delay in installation has occurred as a result of the supplier's dereliction and has exceeded a period of one month from the date of receiving the items, the approved warranty period will automatically be considered as 24 months from the date of installation.

- If at any time during the warranty period the item becomes inoperative due to a technical fault the item must then be repaired by the supplier /local agent within a period of 5 working days from written notification, warranty will be extended according to downtime period.

- If the delay exceeds 30 days the supplier must replace the item with a new identical functioning one (within the same delivery period mentioned in the final offer). In case the item was replaced by a new one, the warranty period mentioned in 9.a) above will start from the installation and commissioning date of the new item.

- Local agent/ supplier is committed to transport and install any awarded item during the warranty period on free of charge basis to any location inside the country whenever required by Royal Medical Services, this should not include pre-installation requirements in the new site.

- All offered items (main unit) should be fully designed, manufactured, and labelled by their real original manufacturer in which all related testing, research, development and approvals went through.

Any relabelled products for the main unit (white-label manufacturing, OEM, or repackaging) are rejected.

- One set of operation manual(s) and one set of service manual(s) including schematics and a spare-part list must be delivered with each unit, CD/DVD is acceptable. For large tenders, a certain agreed percentage of manuals per item may be agreed upon.

- Where applicable, pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include removal of old system(s), any civil work, electrical work or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new system(s).

- Power requirements: where applicable either single phase 220V, 50Hz or 3-phase 380V. Systems with external transformers are considered conforming only if clearly stated in the technical specifications.

- Technical offers must include clear original technical brochures/catalogues for all offered items.

- Offers must include fully detailed technical offers and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s), full technical description/specifications and any accessories or options included in the offer.

Offer must clearly indicate the origin of the offered item and the country where the manufacturer of the offered items is based.

- Compliance sheets must be as per the tabular format of the technical specifications in the tender documents, listing the required specifications on one column and a  Yes  or NO  response to each point in the adjacent column, with reference to page and line numbers in the relevant technical brochure. Offers not complying with this term shall be rejected.

- Qualifications and after sales service: The technical bid should contain all the necessary documents (training certificates) to prove the capability of the bidders for supplying and installing a trouble free equipment meeting the quality standards and technical specifications of the manufacturer and the ability of the bidders for providing efficient after sales service conducted by authorized certified biomedical service engineers with minimum 2 years experience in the same field.

- Accessories and consumables:

  - Any accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately.

All offered accessories and consumables must be approved by the manufacturer.

  - Technical offers must include a priced list for all accessories and consumables related to the required equipment as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) with prices fixed for a period of five years from the date of installation and commissioning with a maximum annual increase of 2%, any essential item not listed will be considered free of charge.

  - Accessories and consumables must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Where applicable, a start-up kit of accessories and consumable items must be provided with each system on a free-of-charge basis.

- Spare Parts:

  - Technical offers must include a comprehensive and priced list of all spare parts related to the awarded equipment (including rechargeable batteries) as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) valid for a minimum period of five years with a maximum annual increase of 2%, commencing at the end date of the warranty period, any essential item not listed will be considered free of charge.

  - Spare parts must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Prices of spare parts should be reasonable and will be taken into consideration during the purchasing process; the purchasing committee has the right to eliminate any offered item with unreasonable high prices of spare parts.

  - Delivery period of required spare parts should not exceed 2 months from the date of the final order.

- Spare parts, consumables and accessories availability must be guaranteed for a minimum period of ten years starting from the date of installation and commissioning.

- Local agent is committed to sign a service contract with Royal Medical Services to perform all relevant maintenance for the awarded equipment (whenever required by DRMS within a period that does not exceed one year from the end date of the warranty period and if mentioned clearly in the specifications) without exceeding the original mentioned cost and as per mentioned terms within minimum required period.

- Tender Awards:

  - For the final list of offers having a chance of winning the award, the awarding process shall be based on the accumulative value of both the offered item and its’ running cost (Total Cost of Ownership) over a period of seven years from the date of installation and commissioning. Only offers with the lowest total cost of ownership over a period of seven years from the date of installation and commissioning shall qualify for the award.

  - Running cost includes the value of consumables, accessories needed to operate the system over the same period as well as the cost of any service contract (where applicable).

- For PC/Laptop based systems:

  - Complete restoration medium (CD/DVD/etc.) of the operating system and the application software must be supplied.

  - Purchased licenses, software keys & dongles must be provided at any time during and after the warranty period on free of charge basis whenever required by DRMS for a minimum period of ten years starting from the date of installation and commissioning.

  - Where locally supplied computers, laptops & printers are offered, the offered model should be from well known manufacturer.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores, installation, pre-installation (if needed), training, commissioning, warranty and bringing the equipment into service.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- a. DRMS has the right to increase the awarded quantities by a percentage not exceeding 35% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

b. DRMS has the right to decrease the awarded quantities by a percentage not exceeding 50% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- a. Bidder is not allowed to submit more than one offer for the same item whether that’s solely or in coalition or partnership with other bidders.

b. Bidder is allowed to include within their offer optional alternatives for the same offered item from the same manufacturer.

- The supplier must furnish DRMS with a guarantee stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelve months from the date of final acceptance of the equipment by DRMS.

- Training:

  - For items where abroad service training courses are required in technical specifications, offers must include a certified service training program for at least 3 working days at a reputable center abroad recognized by the manufacturer for at least one biomedical engineer or biomedical technician; all costs inclusive, air tickets, , boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - For items where abroad user  training courses are required in technical specifications, offers must include a certified operator training program at a reputable center abroad recognized by the manufacturer for at least one operator; all costs inclusive, air tickets, boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - The period of the training courses must be according to the manufacturer’s program excluding traveling days and must be stated clearly in the technical offer.

  - Training Programs must conform to the following standards:

- User training must comprise understanding and use of operation manual(s), correct and safe operation of the equipment, as well as user preventive maintenance and calibration.

- Service training must comprise: theory, understanding and use of service manual(s), calibration, preventive maintenance procedure, and practical troubleshooting and repair exercises, and must be conducted by professional instructors employed or authorized by the system manufacturer.

- Service training must be conducted on a system of identical make, model, and configuration to that purchased by DRMS, and designated by the manufacturer or the local agent for training purposes.

- Certificates must be endorsed and officially sealed by the system manufacturer, legally empowering trainees to engage in user and service activities according to operation and service manual(s).

- Where applicable, offers must include an on-site user and service training.

30. For offers submitted in Jordanian dinars, Payment will be either by wire transfer or by cheque after final acceptance of goods. Any other way of payment will be rejected.

Attachment 1

| # | Equipment |
| --- | --- |
| 1 | Adjustable Auto Pipettes |
| 2 | Resuscitation Bag |
| 3 | Laryngoscope Set |
| 4 | Oxygen Flow meter wall type/ single |
| 5 | Regulator Suction with canister, wall vacuum outlet |
| 6 | Oxygen Regulator for Oxygen Cylinders |
| 7 | Pulse Oximeter, Finger type |
| 8 | Oxygen Cylinder |
| 9 | Doppler, portable |
| 10 | Diagnostic set, Portable |
| 11 | Direct Ophthalmoscope, Portable |
| 12 | Otoscope, Portable |
| 13 | Air Mattress System, homecare |
| 14 | Stethoscope |
| 15 | Aneroid Sphygmomanometer |
| 16 | Video Assisted Laryngoscope, portable |
| 17 | Scale, Manual |
| 18 | Wood's Light |
| 19 | Cough Pressure, Normal Saline |
| 20 | Rehabilitation Walking Parallel Bars, non-powered |
| 21 | Therapy Mat |
| 22 | Medical Ball All Size |
| 23 | Dumbbells Rack with complete set of dumbbells |
| 24 | Crutches |
| 25 | Shoulder wheel |
| 26 | Mobile Mirror |
| 27 | Cuff Weights |
| 28 | Walker, different sizes |
| 29 | Patient Elbow Stick |
| 30 | Rehabilitation Training Ladder |
| 31 | Rehabilitation Suspension Frame |
| 32 | Exercise Band All Size (Theraband) |
| 33 | Lens trial set |
| 34 | Wheel Chair |
| 35 | Commode Chair |
| 36 | Bassinet (Baby Cot) |
| 37 | Resuscitation Cart (Crash Cart) |
| 38 | Medication Cart |
| 39 | Cart, Drawers |
| 40 | Examination Couch, Manual |
| 41 | Gynaecology Examination Table, Manual |
| 42 | Examination Table, Neonates, Manual |
| 43 | Intravenous Pole, Mobile Stand |
| 44 | DDA Cabinet |
| 45 | Stainless steel Multipurpose Trolley |
| 46 | Cabinet, Instrument, Operation Theatres |
| 47 | Dressing Cart |
| 48 | Stainless steel wire shelving unit |
| 49 | Paper Trolley |
| 50 | Stainless Steel Sink (Clean up counter) |
| 51 | Scopes Cabinet |
| 52 | Mayo Table |
| 53 | Table, Instrument |
| 54 | Stool, Adjustable, Doctor |
| 55 | Stool, Adjustable, Operation Theatres |
| 56 | Carts, linen/laundry, soiled, Double |
| 57 | Closed distribution trolley |
| 58 | Step Ladder, Conductive, Double |
| 59 | Step, Surgeon, Single |
| 60 | Kick Bucket |
| 61 | Mobile Stand for Oxygen Cylinder |
| 62 | Stainless Steel Wire Basket, 1 STU |
| 63 | Cart, Plaster |

| Annex 6: Medical Furniture Equipment |  |  |
| --- | --- | --- |
| Item 1 | Mayo Instrument Table | Qty. (34) |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design; should be well constructed to withstand typical abuse and cleaning |  |  |
| 2 | Table top, support column and foot pedal should be constructed from heavy-duty stainless steel |  |  |
| 3 | T-shape base that is constructed from heavy duty, shock resistant material that is easy to clean |  |  |
| 4 | Height adjustable by of foot operated hydraulic pump from 950 up to 1200 mm or more with damping mechanism when lowering the table |  |  |
| 5 | Removable stainless steel tray with raised edges approx. 700 x 450 mm. |  |  |
| 6 | Height and swivel blocking system |  |  |
| 7 | Rounded corners and edges |  |  |
| 8 | Double antistatic medical grade swivel castors with diameter of 75 mm or more |  |  |
| 9 | Safe working load: 25 kg or more |  |  |

| Item 2 | Table, Instrument | Qty. (20) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Design: As per figure |  |  |
| 2 | Ergonomic heavy duty design that is suitable to transport heavy surgical instruments (up to 150 kg) inside the OR department |  |  |
| 3 | It should be constructed from heavy duty medical grade stainless steel type 304 or better |  |  |
| 4 | It should include two stainless steel solid shelves, solid and completely open without raised edges |  |  |
| 5 | Without push handles |  |  |
| 6 | All-round bumper or 4 corner bumpers |  |  |
| 7 | It should be easy to clean and disinfect |  |  |
| 8 | Antistatic castors with diameter of around 7.5 cm (2 with brakes) |  |  |
| 9 | Approximate dimensions (L*W*H): 90*50*100 cm |  |  |

| Item 3 | Stainless steel Multipurpose Trolley, 100 x 60 cm | Qty. (32) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design |  |  |
| 2 | Entirely made from high quality stainless steel (304 or better) |  |  |
| 3 | Four strong legs with heavy duty swivel castors. |  |  |
| 4 | Two shelves with upper edges. |  |  |
| 5 | No sharp edges allowed. |  |  |
| 6 | Push handle. |  |  |
| 7 | Dimensions for each shelf ≈ 100 x 60 cm (L*W). |  |  |

| Item 4 | Stool, Adjustable, Operation Theatres | Qty. (64) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Rolling swivel stool with round seat and backrest, heavy duty design |  |  |
| 2 | # The stool main frame should be constructed from stainless steel material with base and castors from an equivalent durable material |  |  |
| 3 | The stool seat should be circular with 360 degree swivel and of minimum 15” diameter and made from high density foam with manually adjustable backrest. |  |  |
| 4 | # The seating should be hygienic and easy to clean |  |  |
| 5 | # Height adjustment from 500–700 mm approximately, foot operated |  |  |
| 6 | # It shall be mounted on five double swivel castors of 50 mm approximate diameter, at least two of them should be electrically conductive |  |  |
| 7 | # The stool should be without hand rests |  |  |
| 8 | # Loading capacity at least 120 kg |  |  |
| 9 | The stool should have full circle heel rest. |  |  |

| Item 5 | Closed Distribution Trolley | Qty. (8) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design |  |  |
| 2 | Entirely made from heavy duty stainless steel. |  |  |
| 3 | Four heavy duty swivel castors. |  |  |
| 4 | Protective pumpers. |  |  |
| 5 | Two inner shelves. |  |  |
| 6 | Two hinged doors with heavy duty securing locking handle. |  |  |
| 7 | No sharp edges allowed. |  |  |
| 8 | Push handle. |  |  |
| 9 | Dimensions for each shelf ≈ 60 x 90 cm. |  |  |
| 10 | Height ≈ 100 cm |  |  |

| Item 6 | Stainless Steel Wire Shelving Unit | Qty. (92) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | ## Made entirely in stainless steel |  |  |
| 2 | ## 5 wire shelves. |  |  |
| 3 | ## Heavy duty construction |  |  |
| 4 | ## Approx. dimensions: 100*60*200 cm (L*D*H) each shelf should fit at least 3 standard baskets (30*30*60 cm). |  |  |
| 5 | ## Both models should be quoted separately; mobile and fixed |  |  |

| Item 7 | Intravenous Pole, Mobile stand | Qty. (61) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design |  |  |
| 2 | Constructed entirely from high-quality stainless steel. |  |  |
| 3 | Equipped with at least four Ram’s Horn style hooks. |  |  |
| 4 | Height adjustable by means of a telescoping upright rod. |  |  |
| 5 | The selected height to be securely fixed via a twist lock, knob handle |  |  |
| 6 | Mounted 5 antistatic anti-scratch swivel-type castors. |  |  |

| Item 8 | Kick Bucket | Qty. (64) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Constructed entirely from high quality stainless steel. |  |  |
| 2 | Mounted on at least 4 castors |  |  |
| 3 | Single bucket. |  |  |
| 4 | Minimum capacity 10 litres |  |  |

| Item 9 | Resuscitation Cart (Crash Cart) | Qty. (16) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design |  |  |
| 2 | This cart should be an advanced design that manoeuvrable, tightly organized with easy access to all of its parts. |  |  |
| 3 | Heavy duty design |  |  |
| 4 | It should be constructed from polymer material |  |  |
| 5 | It should be mounted on full-swivel casters with brakes with at least 10 cm diameter |  |  |
| 6 | Ergonomic handle |  |  |
| 7 | It should have a quick release master security latch/lock mechanism with plastic seals that secures the entire cart |  |  |
| 8 | All-round or corner bumpers |  |  |
| 9 | It should include at least 5 drawers as following: |  |  |
|  | - One small size medication drawer (3” height) with dividers |  |  |
|  | - Three (or more) multipurpose medium size drawers (6” height) with dividers. |  |  |
|  | - One (or more) large drawer (9” height) with dividers. |  |  |
| 10 | It should have the following features and accessories as a minimum: |  |  |
|  | - Defibrillator shelf for mounting Defibrillator/Monitor |  |  |
|  | - Oxygen tank holder with oxygen cylinder with regulator |  |  |
|  | - Sharps container with holder |  |  |
|  | - Waste bin with holder |  |  |
|  | - Back board mounted to the front or back of cart |  |  |
|  | - One IV pole |  |  |

| Item 10 | Cart, Anaesthesia | Qty. (16) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design |  |  |
| 2 | It should be constructed from corrosion proof, heavy duty, lightweight polymer material |  |  |
| 3 | It shall be disinfectant proof and easy to clean |  |  |
| 4 | The cart shall be mobile on large swivel casters of at least 10 cm height with locks |  |  |
| 5 | ## Wrap around or corner bumpers |  |  |
| 6 | At least 18 large storage space bins with label holder, dimensions of drawers should be clearly indicated |  |  |
| 7 | ## Two lower drawers, 6 inches height |  |  |
| 8 | ## Master lock or locking mechanism that secures the entire cart |  |  |
| 9 | ## Retractable side shelf |  |  |
| 10 | ## Upper accessory rail (overbridge) with at least (8) tilt bins |  |  |
| 11 | ## It should include a work upper surface with raised edges |  |  |
| 12 | ## Push handle |  |  |
| 13 | ## IV pole holder with IV pole having double hooks |  |  |

| Item 11 | Medical/ Surgical Ward Bed, Electrical | Qty. (18) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | Durable & heavy duty bed from a reputable well known manufacturer |  |  |
| A | Patient bed |  |  |
|  | - Heavy-duty construction suitable for rugged use |  |  |
|  | - Safe working load of the bed to be at least 220 Kg. |  |  |
|  | - The dimensions of the sleeping surface platform to be minimum of (87*200) cm. |  |  |
|  | - Bed extension |  |  |
|  | - Mattress for bed extension to be prices separately (for 20 beds only) |  |  |
|  | - Flat sleep surface made from heavy duty material, at least head section should be radiolucent |  |  |
|  | - The bed must comply with the international standard IEC/ EN 60601-2-52 |  |  |
|  | - Configuration: Four sections platform: - Backrest (head of bed) section - Seat section - Thigh section - Foot (calf section) |  |  |
|  | - The bed should be fully electrically operated; the following manoeuvres must be included: - Height adjustment of the bed: ≤ 40 cm to ≥ 75 cm. (measured without mattress from floor to the top of the platform) - Backrest adjustable: 0° to ≥60° - Trendelenburg/ Reverse trendelenburg: +12°/-12° or more |  |  |
|  | - Height adjustment technique: Scissors type |  |  |
|  | - Clearance between dropped-down side rails and the floor on bed's lowest position (under-bed clearance) should be at least 10 cm |  |  |
|  | - Easy released manual CPR, it should be fully accessible even when side rails are in lowest position. |  |  |
|  | - Backrest (head of bead) angle display |  |  |
|  | - Left and right side rails: - Split and heavy duty model - With embedded integrated caregiver control panels on both sides of the bed (upper left and upper right) on the outer sides of the rails. - At least one embedded patient control panel at the inner side of the side rail. - The side rails can be manually raised up and dropped down with damping and top locking mechanism. - With safe gap concept that prevents patient head or limbs entrapment (free of hazardous gaps). - Side rails must cover at least ⅔ of the sleeping platform |  |  |
|  | - Heavy duty non-metallic Headboard and Footboard, at least headboard should be removable. |  |  |
|  | - Four heavy-duty medical grade electro-conductive, non-marking castors of not less than 15 cm diameter. |  |  |
|  | - Central brake system with steering function |  |  |
|  | - Four original protective revolving rubber bumpers: Two of which should be located at the corners of the bed end, the other two should be located at the head of the bed. |  |  |
|  | - At least two IV-pole sockets |  |  |
|  | - One compatible IV pole with double hooks |  |  |
|  | - Drainage bag hook located at each side of the bed |  |  |
|  | - Sleeping surfaces: - Must cover the whole sleeping surface platform - Pressure ulcer preventive - Foam core - At least 14 cm thickness - Maximum patient weight limit ≥ 150 kg - Fire retardant, anti-microbial and anti-fungus - Complete with non-latex, fluid resistant and washable cover, welded or with concealed zip |  |  |
|  | - All available options and accessories must be offered and priced separately. |  |  |
|  | - A list of standard accessories for the offered model |  |  |
| B | ## Bedside Cabinet |  |  |
|  | - Standalone. |  |  |
|  | - Heavy-duty and robust. |  |  |
|  | - Complete with a minimum of one drawer. |  |  |
|  | - Lower cabinet. |  |  |
|  | - Mounted on four heavy-duty castors. |  |  |
|  | - Material, colour and design must be compatible with the beds and all available colours to be submitted with the offered. |  |  |
| C | ## Over-bed table. |  |  |
|  | - Standalone |  |  |
|  | - Heavy-duty and robust. |  |  |
|  | - Mounted on four heavy-duty castors. |  |  |
|  | - Material, colour and design must be compatible with the beds and all available colours to be submitted with the offered. |  |  |
| D | All three components (patient bed, bed side cabinet and over bed table should be from the same manufacturer. |  |  |

| Item 12 | Transport Stretcher, Mobile, Hospital | Qty. (46) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The stretcher shall be used as a Transport Stretcher to transfer patients between different hospital departments. It shall provide a high level of comfort and safety for the patient while providing the caregiver with an advanced patient transport system having advanced mobility features. |  |  |
| 2 | From a well-known manufacturer, heavy duty design; the stretcher should be well constructed to withstand typical abuse and cleaning |  |  |
| 3 | Dual column design |  |  |
| 4 | Height adjustable via a foot operated hydraulic pump at each side of the trolley |  |  |
| 5 | Backrest and trendelenburg adjustable via pneumatic cylinders. |  |  |
| 6 | Overall length: 210 cm |  |  |
| 7 | Overall width: min 75 cm. (side rails down) |  |  |
| 8 | Patient surface should be at least: 65 * 190 cm |  |  |
| 9 | Safe working load: ≥ 300 Kg. |  |  |
| 10 | It shall have controls for the following: |  |  |
|  | - Height range adjustment (foot operated): Low ≤ 60 cm High ≥ 85 cm. |  |  |
|  | - Backrest: 0° to 90°. |  |  |
|  | - Trendelenburg and reverse Trendelenburg (hands-free control): ≥ ±15°. |  |  |
| 11 | It shall include an 8 cm thick patient mattress that is conductive and flame resistant. |  |  |
| 12 | Push handles. |  |  |
| 13 | Foldable side rails. |  |  |
| 14 | Rubber bumpers |  |  |
| 15 | Four integrated IV receptacles. |  |  |
| 16 | One IV pole with dual IV hooks. |  |  |
| 17 | ## Integrated oxygen bottle holder |  |  |
| 18 | Large (min 20 cm diameter) swivel/steerable medical grade antistatic castors with brakes. |  |  |
| 19 | Central brake system |  |  |
| 20 | A list of standard accessories for the offered model |  |  |

| Item 13 | Wheel Chair, Adult | Qty. (8) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | Heavy duty design |  |  |
| A | # Wheels: |  |  |
|  | - The user propelled wheels should be 24 inches in diameter mounted into the rear of both side frames and should have a push rim attached with toggle or lever wheel locks |  |  |
|  | - Swivel casters (five to eight inches in diameter) should be mounted to the front of both side frames. |  |  |
|  | - Wheels shall be tubeless or air free. |  |  |
| B | Seat and Backrest: |  |  |
|  | - The wheelchair should have collapsible seat and backrest. |  |  |
|  | - The wheelchair should have some type of armrest, and these armrests shall be adjustable in height and pivoting or removable for ease of transfers. |  |  |
|  | - The wheelchair should have leg rests or footrests, and they shall have adjustable lengths and either detachable or swing-away features for ease of transfer. Footrests should either flip-up or fold and should have impact guards. |  |  |
|  | - The wheelchair should have anti-tip devices and should not tip on inclines with slopes of up to 15°. |  |  |
|  | - The wheelchair upholstery should pass applicable standards for flame resistance. |  |  |
|  | - Approximate overall dimensions (w x L x h): 65 x 95 x 80 cm. |  |  |

| Item 14 | Wheel Chair, Paediatric | Qty. (6) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| * | Heavy duty design |  |  |
| A | # Wheels: |  |  |
|  | - The user propelled wheels should be 24 inches in diameter mounted into the rear of both side frames and should have a push rim attached with toggle or lever wheel locks |  |  |
|  | - Swivel casters (five to eight inches in diameter) should be mounted to the front of both side frames. |  |  |
|  | - Wheels shall be tubeless or air free. |  |  |
| B | Seat and Backrest: |  |  |
|  | - The wheelchair should have collapsible seat and backrest. |  |  |
|  | - The wheelchair should have some type of armrest, and these armrests shall be adjustable in height and pivoting or removable for ease of transfers. |  |  |
|  | - The wheelchair should have leg rests or footrests, and they shall have adjustable lengths and either detachable or swing-away features for ease of transfer. Footrests should either flip-up or fold and should have impact guards. |  |  |
|  | - The wheelchair should have anti-tip devices and should not tip on inclines with slopes of up to 15°. |  |  |
|  | - The wheelchair upholstery should pass applicable standards for flame resistance. |  |  |
|  | - Approximate overall dimensions (w x L x h): 50 x 75 x 80 cm. |  |  |

| Item 15 | Medication Cart | Qty. (6) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design |  |  |
| 2 | The medication cart shall be a unit-dose cart with individual bins or drawers for each patient’s dosage. |  |  |
| 3 | Ergonomic heavy duty design |  |  |
| 4 | It should be constructed from polymer material |  |  |
| 5 | It should include a work upper surface. |  |  |
| 6 | At least 20 large storage space bins of approximate size 20 cm*10 cm*14 cm (L*W*H) |  |  |
| 7 | Large lower utility drawer with removable adjustable dividers (3 longitudinal sections) |  |  |
| 8 | Master lock or locking mechanism that secures the entire cart and that must be unlocked before any individual compartment can be opened |  |  |
| 9 | A push handle. |  |  |
| 10 | All-round or corner bumpers |  |  |
| 11 | Waste container |  |  |
| 12 | The cart shall be mobile on large swivel casters with locks |  |  |

| Item 16 | DDA Cabinet, Double Door, Wards | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # The unit should be constructed from heavy duty epoxy powder coated steel with enamelled finish & rounded corners |  |  |
| 2 | ## The unit should have double doors (inner and outer) with a secure independent and different lock for each door |  |  |
| 3 | ## The unit should have at least two shelves. |  |  |
| 4 | ## The unit should be securely mounted on wall. |  |  |
| 5 | ## No sharp edges |  |  |
| 6 | ## Dimensions: 60*50*35 cm (H*W*D) approximately |  |  |

| Item 17 | Medication Cabinet (Non-Narcotics), Single Door, Wall mounted | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # The unit should be constructed from heavy duty epoxy powder coated steel with enamelled finish & rounded corners |  |  |
| 2 | ## The unit should have single door with a secure lock |  |  |
| 3 | ## At least two shelves. |  |  |
| 4 | ## The unit should be securely mounted on wall. |  |  |
| 5 | ## No sharp edges |  |  |
| 6 | ## Dimensions: 60*50*35 cm (H*W*D) approximately |  |  |

| Item 18 | Dressing Cart, 900 X 600 | Qty. (10) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Heavy duty design |  |  |
| 2 | Should be designed to transport dressing supplies and instruments |  |  |
| 3 | It shall not have side panels. |  |  |
| 4 | Approximate overall dimensions approximately (W x D x H): 900 x 600 x 900 mm with 4-leg frame construction. |  |  |
| 5 | It shall be made completely of medical grade stainless steel 304 or better with high quality finish to assure durability. |  |  |
| 6 | Removable waste container |  |  |
| 7 | The cart should have two shelves, the upper shelf should guard rails on 3 sides |  |  |
| 8 | One large utility drawer under the top shelf |  |  |
| 9 | It shall be mounted on 4 castors with brakes (at least 2) |  |  |

| Item 19 | Mobile Stand for Oxygen Cylinder | Qty. (14) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | # Heavy duty design |  |  |
| 2 | Stand with two castors. |  |  |
| 3 | Push / pull handle. |  |  |
| 4 | Variable size stands to suit several capacity cylinders should be quoted (1,2 & 7 cubic meters cylinders) |  |  |
| 5 | Securing mechanism |  |  |

| Item 20 | Cart, Drawers | Qty. (3) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Ergonomic heavy duty design |  |  |
| 2 | It should be constructed from corrosion proof polymer material |  |  |
| 3 | Supply cart with 5 drawers: - Three small drawers with dividers - One Medium size drawer - Large lower utility drawer |  |  |
| 4 | A push handle. |  |  |
| 5 | Master lock or locking mechanism |  |  |
| 6 | The cart shall be mobile on large swivel casters with locks |  |  |
| 7 | ## Wrap around bumper |  |  |
| 8 | Waste container |  |  |
| 9 | All-round or corner bumpers |  |  |

SPECIAL TERMS

- Offers not complying with any of the special terms or the technical specifications shall be considered non-conforming with tender requirements.

- Any vendor providing FORGED documents shall be disqualified from the current tender and banned from participating in any future RMS tenders.

- Wherever term “based” is mentioned it refers to the country where the manufacturing company is founded & established.

- All equipment must be the most recently released model/version which is equal to or higher than the range of the specifications of the required system (low, mid or high) and equal to or higher than the level of technology and required options mentioned in  the technical specifications.

- Required certificates (must be submitted with the technical offer):

- FDA clearance (510K) for equipment of USA based companies.

- MDR (EU) 2017/745 certificate for equipment of EU (European Union) based companies.

- UKCA certificate For Equipment of Great Britain based companies (England, Scotland, Northern Ireland, and Wales).

- ARTG (Australian Register of Therapeutic Goods) certification or approval for Australian and New Zealand based companies.

- Establishment License with the relevant Device License issued by the Therapeutic Products Directorate for Equipment of Canadian based companies.

- PMDA (Pharmaceuticals and Medical Devices Agency) certification or approval for equipment of Japanese based companies.

- Swissmedic (Swiss agency for therapeutic products) certification or approval for equipment of Swiss based companies.

- Norwegian Medicines Agency certification or approval for equipment of Norwegian based companies.

- Only for class Ι medical equipment manufactured by companies based in one of the countries mentioned above, submission of either one of the certificates mentioned above or a free sale certificate in any of these countries shall be accepted.

- For the following equipment:

- All X-ray equipment, MRI, ultrasound, and nuclear medicine systems (regardless where the manufacturing companies are based).

- Where the manufacturing companies are based in other origins than the mentioned in terms 2.1 - 2.8.

The following are required:

- At least two of the certificates mentioned above, one of which has to be FDA clearance (510K) (Only for class Ι medical equipment submission of certificates mentioned in 2.9 shall be accepted).

- Evaluation certificate from the Royal Medical Services for the same offered model with at least 80% passing grade.

If the evaluation is not applicable (based on purchasing committee perspective) bidder should submit a list of installation basis of the same offered model and/ or previous models in at least two of the following hospitals (King Hussein Cancer Center, National Center for Diabetes Endocrinology and Genetic Diseases, Jordan University Hospital or King Abdullah University Hospital) with at least three years of operation, list should include: Name of hospital, Model installed, Quantity, and date of installation.

The purchasing committee has the right to officially contact any of these hospitals and disqualify any offer where the feedback is negative in operation, after sales service or local agent performance.

- The vendor is responsible to ensure through official  documents that classified medical devices are manufactured in conformity with applicable quality system standards (ISO, IEC); (the international quality systems standard for medical device manufacturing ISO 13485) or ISO9001.

- With each offer, bidders must provide a formally endorsed document issued by the manufacturer stating that the bidder is the sole certified agent for the offered item.

- In all of the above cases (except 2.12) certificates must be formally endorsed by JFDA.

- Any vendor not submitting all required certificates will be eliminated.

- Official website of the manufacturer must be provided. The website must clearly demonstrate the history profile and manufacturing features of the company including the same offered model and its brochure.

In case of inconsistency between the information on the website with the provided information of the local agent regarding any offered product, the committee has the right to eliminate that product from the awarding process.

- Offered items should be from reputable well known manufacturers and excellent experience in the field and shall have multiple installations of the same offered model and/ or previous models in RMS main hospitals with at least two years of operation and excellent experience in operation, after sales service & local agent performance; otherwise the purchasing committee has the right to request any of the following:

- An evaluation certificate as mentioned in term 2.10.b.

- A sample of the same offered model at any time during the purchasing process, the sample unit shall be delivered within three weeks from the date of a written notification; any vendor rejects to deliver a sample unit will be eliminated. Evaluation duration will be determined by the purchasing committee.

Any offered item fail in the evaluation/assessment process will be rejected

- Vendors must specify the origin of the offered items and accessories in the technical offer.

- Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first half of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period shall be rejected.

- Goods are to be brand new; manufacturing date of the delivered products should not exceed 18 months from the date of the final award.

- Delivery period should be mentioned clearly; the purchasing committee has the right to reject any offer with delivery time exceeding 4 months.

- Warranty:

- Offers must include a full warranty for a period of a minimum of 24 months from the date of installation or 30 months from the date of receiving the items at the agreed location mentioned in the final order (whichever comes first).

Warranty must include corrective & preventive maintenance activities as per manufacturer recommendations including:

- Required spare parts (free of charge)

- Labour

- Hardware

- Software

- Rechargeable batteries

At the end of the warranty period, The supplier commits to implement final inspection of the submitted goods and submit the reports signed by the site chief engineer stating that the equipment are working properly as well as all preventive maintenance reports during the warranty period.

Warranty & installation are excluded for items mentioned in Attachment no.1 the purchasing committee has the right to exclude any other item not listed in the attachment and does not required installation and warranty.

- In the case where a delay in installation has occurred as a result of the supplier's dereliction and has exceeded a period of one month from the date of receiving the items, the approved warranty period will automatically be considered as 24 months from the date of installation.

- If at any time during the warranty period the item becomes inoperative due to a technical fault the item must then be repaired by the supplier /local agent within a period of 5 working days from written notification, warranty will be extended according to downtime period.

- If the delay exceeds 30 days the supplier must replace the item with a new identical functioning one (within the same delivery period mentioned in the final offer). In case the item was replaced by a new one, the warranty period mentioned in 9.a) above will start from the installation and commissioning date of the new item.

- Local agent/ supplier is committed to transport and install any awarded item during the warranty period on free of charge basis to any location inside the country whenever required by Royal Medical Services, this should not include pre-installation requirements in the new site.

- All offered items (main unit) should be fully designed, manufactured, and labelled by their real original manufacturer in which all related testing, research, development and approvals went through.

Any relabelled products for the main unit (white-label manufacturing, OEM, or repackaging) are rejected.

- One set of operation manual(s) and one set of service manual(s) including schematics and a spare-part list must be delivered with each unit, CD/DVD is acceptable. For large tenders, a certain agreed percentage of manuals per item may be agreed upon.

- Where applicable, pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include removal of old system(s), any civil work, electrical work or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new system(s).

- Power requirements: where applicable either single phase 220V, 50Hz or 3-phase 380V. Systems with external transformers are considered conforming only if clearly stated in the technical specifications.

- Technical offers must include clear original technical brochures/catalogues for all offered items.

- Offers must include fully detailed technical offers and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s), full technical description/specifications and any accessories or options included in the offer.

Offer must clearly indicate the origin of the offered item and the country where the manufacturer of the offered items is based.

- Compliance sheets must be as per the tabular format of the technical specifications in the tender documents, listing the required specifications on one column and a  Yes  or NO  response to each point in the adjacent column, with reference to page and line numbers in the relevant technical brochure. Offers not complying with this term shall be rejected.

- Qualifications and after sales service: The technical bid should contain all the necessary documents (training certificates) to prove the capability of the bidders for supplying and installing a trouble free equipment meeting the quality standards and technical specifications of the manufacturer and the ability of the bidders for providing efficient after sales service conducted by authorized certified biomedical service engineers with minimum 2 years experience in the same field.

- Accessories and consumables:

  - Any accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately.

All offered accessories and consumables must be approved by the manufacturer.

  - Technical offers must include a priced list for all accessories and consumables related to the required equipment as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) with prices fixed for a period of five years from the date of installation and commissioning with a maximum annual increase of 2%, any essential item not listed will be considered free of charge.

  - Accessories and consumables must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Where applicable, a start-up kit of accessories and consumable items must be provided with each system on a free-of-charge basis.

- Spare Parts:

  - Technical offers must include a comprehensive and priced list of all spare parts related to the awarded equipment (including rechargeable batteries) as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) valid for a minimum period of five years with a maximum annual increase of 2%, commencing at the end date of the warranty period, any essential item not listed will be considered free of charge.

  - Spare parts must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Prices of spare parts should be reasonable and will be taken into consideration during the purchasing process; the purchasing committee has the right to eliminate any offered item with unreasonable high prices of spare parts.

  - Delivery period of required spare parts should not exceed 2 months from the date of the final order.

- Spare parts, consumables and accessories availability must be guaranteed for a minimum period of ten years starting from the date of installation and commissioning.

- Local agent is committed to sign a service contract with Royal Medical Services to perform all relevant maintenance for the awarded equipment (whenever required by DRMS within a period that does not exceed one year from the end date of the warranty period and if mentioned clearly in the specifications) without exceeding the original mentioned cost and as per mentioned terms within minimum required period.

- Tender Awards:

  - For the final list of offers having a chance of winning the award, the awarding process shall be based on the accumulative value of both the offered item and its’ running cost (Total Cost of Ownership) over a period of seven years from the date of installation and commissioning. Only offers with the lowest total cost of ownership over a period of seven years from the date of installation and commissioning shall qualify for the award.

  - Running cost includes the value of consumables, accessories needed to operate the system over the same period as well as the cost of any service contract (where applicable).

- For PC/Laptop based systems:

  - Complete restoration medium (CD/DVD/etc.) of the operating system and the application software must be supplied.

  - Purchased licenses, software keys & dongles must be provided at any time during and after the warranty period on free of charge basis whenever required by DRMS for a minimum period of ten years starting from the date of installation and commissioning.

  - Where locally supplied computers, laptops & printers are offered, the offered model should be from well known manufacturer.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores, installation, pre-installation (if needed), training, commissioning, warranty and bringing the equipment into service.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- a. DRMS has the right to increase the awarded quantities by a percentage not exceeding 35% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

b. DRMS has the right to decrease the awarded quantities by a percentage not exceeding 50% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- a. Bidder is not allowed to submit more than one offer for the same item whether that’s solely or in coalition or partnership with other bidders.

b. Bidder is allowed to include within their offer optional alternatives for the same offered item from the same manufacturer.

- The supplier must furnish DRMS with a guarantee stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelve months from the date of final acceptance of the equipment by DRMS.

- Training:

  - For items where abroad service training courses are required in technical specifications, offers must include a certified service training program for at least 3 working days at a reputable center abroad recognized by the manufacturer for at least one biomedical engineer or biomedical technician; all costs inclusive, air tickets, , boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - For items where abroad user  training courses are required in technical specifications, offers must include a certified operator training program at a reputable center abroad recognized by the manufacturer for at least one operator; all costs inclusive, air tickets, boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - The period of the training courses must be according to the manufacturer’s program excluding traveling days and must be stated clearly in the technical offer.

  - Training Programs must conform to the following standards:

- User training must comprise understanding and use of operation manual(s), correct and safe operation of the equipment, as well as user preventive maintenance and calibration.

- Service training must comprise: theory, understanding and use of service manual(s), calibration, preventive maintenance procedure, and practical troubleshooting and repair exercises, and must be conducted by professional instructors employed or authorized by the system manufacturer.

- Service training must be conducted on a system of identical make, model, and configuration to that purchased by DRMS, and designated by the manufacturer or the local agent for training purposes.

- Certificates must be endorsed and officially sealed by the system manufacturer, legally empowering trainees to engage in user and service activities according to operation and service manual(s).

- Where applicable, offers must include an on-site user and service training.

30. For offers submitted in Jordanian dinars, Payment will be either by wire transfer or by cheque after final acceptance of goods. Any other way of payment will be rejected.

Attachment 1

| # | Equipment |
| --- | --- |
| 1 | Adjustable Auto Pipettes |
| 2 | Resuscitation Bag |
| 3 | Laryngoscope Set |
| 4 | Oxygen Flow meter wall type/ single |
| 5 | Regulator Suction with canister, wall vacuum outlet |
| 6 | Oxygen Regulator for Oxygen Cylinders |
| 7 | Pulse Oximeter, Finger type |
| 8 | Oxygen Cylinder |
| 9 | Doppler, portable |
| 10 | Diagnostic set, Portable |
| 11 | Direct Ophthalmoscope, Portable |
| 12 | Otoscope, Portable |
| 13 | Air Mattress System, homecare |
| 14 | Stethoscope |
| 15 | Aneroid Sphygmomanometer |
| 16 | Video Assisted Laryngoscope, portable |
| 17 | Scale, Manual |
| 18 | Wood's Light |
| 19 | Cough Pressure, Normal Saline |
| 20 | Rehabilitation Walking Parallel Bars, non-powered |
| 21 | Therapy Mat |
| 22 | Medical Ball All Size |
| 23 | Dumbbells Rack with complete set of dumbbells |
| 24 | Crutches |
| 25 | Shoulder wheel |
| 26 | Mobile Mirror |
| 27 | Cuff Weights |
| 28 | Walker, different sizes |
| 29 | Patient Elbow Stick |
| 30 | Rehabilitation Training Ladder |
| 31 | Rehabilitation Suspension Frame |
| 32 | Exercise Band All Size (Theraband) |
| 33 | Lens trial set |
| 34 | Wheel Chair |
| 35 | Commode Chair |
| 36 | Bassinet (Baby Cot) |
| 37 | Resuscitation Cart (Crash Cart) |
| 38 | Medication Cart |
| 39 | Cart, Drawers |
| 40 | Examination Couch, Manual |
| 41 | Gynaecology Examination Table, Manual |
| 42 | Examination Table, Neonates, Manual |
| 43 | Intravenous Pole, Mobile Stand |
| 44 | DDA Cabinet |
| 45 | Stainless steel Multipurpose Trolley |
| 46 | Cabinet, Instrument, Operation Theatres |
| 47 | Dressing Cart |
| 48 | Stainless steel wire shelving unit |
| 49 | Paper Trolley |
| 50 | Stainless Steel Sink (Clean up counter) |
| 51 | Scopes Cabinet |
| 52 | Mayo Table |
| 53 | Table, Instrument |
| 54 | Stool, Adjustable, Doctor |
| 55 | Stool, Adjustable, Operation Theatres |
| 56 | Carts, linen/laundry, soiled, Double |
| 57 | Closed distribution trolley |
| 58 | Step Ladder, Conductive, Double |
| 59 | Step, Surgeon, Single |
| 60 | Kick Bucket |
| 61 | Mobile Stand for Oxygen Cylinder |
| 62 | Stainless Steel Wire Basket, 1 STU |
| 63 | Cart, Plaster |

| Annex 8: Laboratory Equipment |  |  |
| --- | --- | --- |
| Item 1 | Fully Automated clinical chemistry Analyser Mid-Range | Qty. (1) |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Brand new Random Access, fully automated floor standing |  |  |
| 2 | A list of standard accessories for the offered model (optional). |  |  |
| 3 | A minimum throughput Not less than 300 tests per hour without ISE continuous loading |  |  |
| 4 | Continuous sample loading. no need to stop the system |  |  |
| 5 | Not less than 25 on-board chemistries |  |  |
| 6 | Ready to use reagents , no need to preparation |  |  |
| 7 | No manual pre-treatment of whole blood samples |  |  |
| 8 | Complete with built-in refrigerated compartment (Cooling unit not external separate unit). |  |  |
| 9 | Sample and reagent level detection |  |  |
| 10 | Wide variety of automatic selection of not less than 12 visible wavelengths: 340-700nm. |  |  |
| 11 | Automatic dilution of abnormal readings |  |  |
| 12 | Instrument to be attached to intra-net interface system |  |  |
| 13 | Real time of reagent/sample tray reaction status |  |  |
| 14 | Stat sample priority not less than 5 samples |  |  |
| 15 | High quality brand name Laser printer (local supplied printers are accepted). |  |  |
| 16 | Computer interface and LCD screen. |  |  |
| 17 | All the running requirements for the system including reagents, consumables, electrodes, full quality control daily, calibration and maintenance kits sufficient as annex. (A). |  |  |
| 18 | Pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include any civil work, R.O. Water system, electrical work and UPS or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new systems if required. |  |  |
| 19 | It should support the connection to an internal LAN system or host computer as part of Hospital Information System or Laboratory Information System |  |  |
| 20 | Bidirectional communication for Hospital Information System or Laboratory Information System |  |  |
| 21 | ISO requirements(reagents, documents, calculation, procedures and reference range) must be included upon request |  |  |

| Item 2 | Ordinary Light Microscope | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | 6-fold revolving objective nosepiece, rear facing. |  |  |
| 3 | Semi-Apochromatic infinity objectives; 4x, 10x, 20x, 40x and (100x). All objective to be of the same class for use with up to 22mm field of view. |  |  |
| 4 | Ergonomic binocular tube, F.O.V. 22mm, |  |  |
| 5 | Pair of adjustable eyepieces 10x/22mm. With IPD adjustment. |  |  |
| 6 | Transmitted light illumination with collector and halogen lamp not less 30W or LED illumination. |  |  |
| 7 | Coaxial double knob focusing and adjustable height stop |  |  |
| 8 | Condenser: Swing-out with aperture iris diaphragm with variable Koehler illumination. |  |  |
| 9 | Mechanical stage with wide scanning area not less than 75 x 50mm with slide holder for up to two slides. |  |  |
| 10 | Simple polarizing set |  |  |
| 11 | Filters: Standard set of filter |  |  |
| 12 | Adjustable light intensity |  |  |
| 13 | All necessary standard accessories and spares to be included. |  |  |
| 14 | A list of additional accessories for the offered model (optional) |  |  |

| Item 3 | Automated Coagulation Analyser (Mid-Range) | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N), Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Automated, walk-away Bench-top computerized blood coagulation analyser for determination of human plasma coagulation time and other parameters using the clot formation method for detection. The offered model must be the highest end. |  |  |
| 2 | Coagulation parameters measurement of at least 8 parameters (PT, PTT, Fibrinogen and D-Dimer). |  |  |
| 3 | Temperature-controlled system. |  |  |
| 4 | Throughput: Not less than 60 test/hour PT/PTT |  |  |
| 5 | Automatic monitoring of reagent levels |  |  |
| 6 | at least 25 on board reagents with 15 in cooling area and 5 stirred position |  |  |
| 7 | Random access sampling with the capability to run a stat sample at any time. |  |  |
| 8 | Programmable clotting assays. |  |  |
| 9 | Reliable, accurate results on compromised samples - icteric, hemolytic and lipemic |  |  |
| 10 | It should have at least one reagent storage position with one pipette well. |  |  |
| 11 | It should have a built in printer or external printer. |  |  |
| 12 | Methodologies : clotting ,chromogenic and immunological assays |  |  |
| 13 | STAT sample ordering at least 5 samples. |  |  |
| 14 | Convenient for all tubes including paediatric. |  |  |
| 15 | It should have a bar-code reader. |  |  |
| 16 | Bidirectional communication for Hospital Information System or Laboratory Information System |  |  |
| 17 | Internal printer or External printer locally supplied as per special terms. |  |  |
| 18 | Start-up kit for all the running requirements for the system including reagents, consumables, piercer, electrodes and maintenance kits sufficient for 5000 samples (PT and PTT) with quality control daily. |  |  |

| Item 4 | Semi-Automated Urine Analyser | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Strip samples. |  |  |
| 3 | Built-in thermal printer or external. |  |  |
| 4 | Throughput not less than approximately 60 strip/ hour. |  |  |
| 5 | To measure glucose, protein, Bilirubin, urobilinogen, pH, specific gravity, ketonebody, occult blood (RBCs)(WBCs leukocytes). |  |  |
| 6 | Data storage: to store not less than the results of approximately 100 patients. |  |  |
| 7 | It should have the ability to be connected to the Hospital Information System and Laboratory Information System. |  |  |
| 8 | A list of standard accessories for the offered model (optional |  |  |

| Item 5 | Haematology Analyser 3-Part Diff | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Automated at least 3-part diff haematology analyser that utilizes whole blood. Not less than 20 parameters including WBC, RBC, Hb, PCV, MCV, MCH, MCHC, and platelets with histograms for WBC, RBC and platelets. |  |  |
| 3 | Histograms to be displayed on the screen in the main unit and printed out. |  |  |
| 4 | Throughput: Not less than 55 samples / hour. |  |  |
| 5 | Cyanide-free reagents for haemoglobin and WBC. |  |  |
| 6 | Sample size not more than 100 µl. |  |  |
| 7 | Touch screen and barcode reader. |  |  |
| 8 | A minimum of 3 quality control files (Low, Normal, and High). |  |  |
| 9 | Capability to display Levy-Jennings graphs on screen. |  |  |
| 10 | Built-in printer using thermal paper |  |  |
| 11 | Memory capacity: Not less than 1000 samples with or without histograms. |  |  |
| 12 | Reagents to be supplied in single components not in one pack. |  |  |
| 13 | It should support the connection to an internal LAN system or host computer as part of Hospital Information System or Laboratory Information System |  |  |
| 14 | Start-up kit for all the running requirements for the system including reagents, consumables, piercer, electrodes and maintenance kits sufficient for 10000 samples and at least one level quality control daily. |  |  |

| Item 6 | Blood Gas Analyser | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Must NOT use one-single pack reagent, long (shelf life and expired date). |  |  |
| 3 | Must not use gas cylinders. |  |  |
| 4 | Standard accessories list. |  |  |
| 5 | Non-cartridge type reagents. |  |  |
| 6 | For measuring pH, PO2, PCO2 and derivative parameters. Electrolytes (Na+, K+, Ca2+, Cl-)/Haematocrit, Metabolites Glu/Lac/Bilirubin |  |  |
| 7 | Co-oximetry (tHb, O2Hb, HHb, COHb, MetHb, SO2) |  |  |
| 8 | Automatic Calculating HCO3, BE, O2Sat, TCO2. |  |  |
| 9 | Maintenance free system |  |  |
| 10 | Fully automatic calibrations |  |  |
| 11 | Fully automatic clean routine |  |  |
| 12 | Built in thermal printer |  |  |
| 13 | To accepts syringe and capillary samples |  |  |
| 14 | Samples throughput not less than 30 samples per hour |  |  |
| 15 | Temperature controls, clear error messages. |  |  |
| 16 | User friendly software |  |  |
| 17 | It should support the connection to an internal LAN system or host computer as part of Hospital Information System or Laboratory Information System. |  |  |
| 18 | A list of standard accessories for the offered model (optional). |  |  |
| 19 | Start-up kit for, All the running requirements for the system including reagents, consumables, electrodes and maintenance kits sufficient for 100 consecutive days at a rate of 20 FULL blood gas samples per day. Requirements must be listed in a clear and itemized manner. |  |  |
| 20 | Any reagents, kits, calibrations, controls, accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately. Prices fixed for a period of five years from the date of installation, any essential item not listed will be considered free of charge. |  |  |

| Item 7 | Bench-Top General Purpose Centrifuge | Qty. (3) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Digital display and microprocessor-controlled operation |  |  |
| 3 | Brushless motor |  |  |
| 4 | Swing-out rotor |  |  |
| 5 | 2 buckets each holds five to eight approximately 16x110 mm collection tubes without adaptors or fittings. |  |  |
| 6 | 2 buckets each holds seven to ten approximately 12x75 mm collection tubes without adaptors or fittings. |  |  |
| 7 | No adaptors or inserts in the buckets |  |  |
| 8 | Automatic lid lock and interlock with open lid indication light |  |  |
| 9 | Built-in digital timer 99 minutes. |  |  |
| 10 | Programmable speed at least 5000 rpm. Or a wider range. |  |  |
| 11 | A list of standard accessories for the offered model (optional). |  |  |

| Item 8 | Refrigerator, Laboratory | Qty. (3) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | The laboratory refrigerator shall store critical products that demand constant temperature such as cultures, media, serum and specimens. |  |  |
| 3 | Uniform cabinet temperature range at least +2ºC to + 8ºC. |  |  |
| 4 | Forced-air cooling with automatic off cycle defrost. |  |  |
| 5 | Digital temperature display with alarm when out of range. |  |  |
| 6 | Heated glass door. |  |  |
| 7 | Automatic condensate evaporator. |  |  |
| 8 | CFC free compressor. |  |  |
| 9 | Stainless steel interior. |  |  |
| 10 | Illuminated interior. |  |  |
| 11 | Door latches with key lock security and self-closing. |  |  |
| 12 | Adjustable shelves. |  |  |
| 13 | Temperature monitoring, chart recording and alarm system when temp. Out of range. |  |  |
| 14 | 4 heavy duty castors, with swivel and brakes and adjustable levelling feet. |  |  |
| 15 | Capacity: not less than 650 Litters . |  |  |
| 16 | Standard accessories list |  |  |

| Item 9 | Blood Bank Refrigerator with Temperature Recorder | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Not less than 5 shelves or drawers. Each drawer divides to chambers ,each chamber suitable for blood bag 500ml |  |  |
| 3 | Vertical 1-2. Double glazes door. |  |  |
| 4 | Volume not less than 600 litres. |  |  |
| 5 | To accommodate not less than 150 blood bag units. |  |  |
| 6 | Equipped with alarm system to maintain inside temp 2-6C. |  |  |
| 7 | A list of standard accessories for the offered model (optional). |  |  |
| 8 | Temperature monitoring, chart recording and alarm system when temp. Out of range. |  |  |
| 9 | 4 heavy duty castors, with swivel and brakes and adjustable levelling feet. |  |  |

| Item 10 | Manual Plasma Extractor | Qty. (4) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Blood Components Extractor, Separation of plasma from centrifuged whole blood units. |  |  |
| 3 | Compatible with all types of blood bags T&T or T&B ,double ,triple and quadruple bags |  |  |

| Item 11 | Centrifuge, Blood Bag | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Refrigerated centrifuge |  |  |
| 3 | Floor standing model. |  |  |
| 4 | Generates not less than 6000g 4500 rpm |  |  |
| 5 | Equipped with; tachometer, thermistor, timer and brake |  |  |
| 6 | At least 4 selectable programs (Programmable speed, temperature and acceleration / deceleration) |  |  |
| 7 | Rotor to accommodate not less than (12) filled (500ml) blood bags |  |  |
| 8 | Brushless drive motor |  |  |
| 9 | Magnetic rotor identification |  |  |
| 10 | Microprocessor controlled operation with touch keypad |  |  |

| Item 12 | Deep Freezer, (-30) | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Temperate up to – 20 Co. |  |  |
| 3 | For Medical Purposes. |  |  |
| 4 | Complete with digital set and actual temperature display and alarm system. |  |  |
| 5 | Capacity not less than 650 liters |  |  |
| 6 | Standard accessories list. |  |  |
| 7 | Temperature monitoring, chart recording and alarm system when temp. Out of range. |  |  |
| 8 | 4 heavy duty castors, with swivel and brakes and adjustable levelling feet. |  |  |
| 9 | Not less than 5 shelves or drawers. |  |  |

| Item 13 | Deep Freezer, (-80) | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Temperate up to – 86 Co. |  |  |
| 3 | For Medical Purposes. |  |  |
| 4 | Complete with digital set and actual temperature display and alarm system. |  |  |
| 5 | Capacity not less than 650 liters |  |  |
| 6 | Standard accessories list. |  |  |
| 7 | Temperature monitoring, chart recording and alarm system when temp. Out of range. |  |  |
| 8 | 4 heavy duty castors, with swivel and brakes and adjustable levelling feet. |  |  |
| 9 | Not less than 5 shelves or drawers. |  |  |

| Item 14 | Water Bath | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Stainless steel tank well insulated. |  |  |
| 2 | Heater and temperature sensors mounted underneath the tank and low water level sensor |  |  |
| 3 | Corrosion resistant and easy to clean with drain system |  |  |
| 4 | Complete with bath cover and handles |  |  |
| 5 | Capacity: not less than 15 litres |  |  |
| 6 | Adjustable temperature from +25 Cº to + 99 Cº |  |  |
| 7 | Water proof control panel and digital display |  |  |
| 8 | Tube racks to hold 12x75mm and 16x100mm tubes, 30-50 tubes, 30-50 tubes/rack |  |  |
| 9 | Temperature to be adjusted in increments of 1°C |  |  |

| Item 15 | Tube Connecting Device | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | Compact design, heavy duty and easy to use |  |  |
| 2 | Designed for sterile connection between all wet and dry PVC tubing used in blood and blood component with high reliability and efficiency |  |  |
| 3 | LCD display |  |  |
| 4 | Tubing size: - Outside diameter: 3.9 - 4.5 mm - Inside diameter: 2.8 - 3.1 mm - Wall thickness: 0.5 – 0.8 mm |  |  |
| 5 | Cutting module construction: stainless steel |  |  |
| 6 | Lid lock mechanism |  |  |
| 7 | Tube holder clamps |  |  |
| 8 | Overheat protection |  |  |
| 9 | All required consumables for 1000 welding processes should be included and quoted separately (Docks, wafers, blades, cassettes, etc…) |  |  |

| Item 16 | Platelets Agitator and Incubator | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Bench top models. |  |  |
| 3 | The instrument should maintain inside temperature between 20-24C and provided with at least one elliptical platelet agitator. |  |  |
| 4 | LCD digital display |  |  |
| 5 | Incubator provided in outlet power supply. |  |  |
| 6 | The agitator completely into incubator side. |  |  |
| 7 | Agitator Capacity not less than 100 Random Bags |  |  |
| 8 | Microprocessor temperature controller with built-in alarm/monitor and chart recorder |  |  |
| 9 | Forced-air circulation maintains chamber uniformity of +/-1ºC |  |  |
| 10 | Not less than 5 Pull-out Drawers with Label holders |  |  |

| Item 17 | Sealer, Tube | Qty. (2) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Capable of single thermal head sealer blood bags tubing. |  |  |
| 3 | Bench top. |  |  |
| 4 | Electrical type heat sealing |  |  |

| Item 18 | Grossing Station, W/Downdraft Extractor | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | 304 or better stainless-steel work surfaces and panels |  |  |
| 3 | Heavy duty disposer (1-2hp) |  |  |
| 4 | Deep sink with cold/hot water faucet. |  |  |
| 5 | Magnetic instrument bar. |  |  |
| 6 | Flex arm magnifier light |  |  |
| 7 | Flex arm light |  |  |
| 8 | Spray hose assembly |  |  |
| 9 | Built-in end table rinse washes down work surfaces quickly |  |  |
| 10 | Hand –free proximity sensing control for cold/hot water faucet, table-rinse, and disposal |  |  |
| 11 | Pull out writing plate and drawer |  |  |
| 12 | Formalin dispensing carboy |  |  |
| 13 | C- Fold paper towel holder |  |  |
| 14 | Fluorescent or LED light for lighting work station |  |  |
| 15 | Dissecting board |  |  |
| 16 | Dimension not less than ( L140/H200/W70) cm. |  |  |
| 17 | Adjustable levelling feet |  |  |
| 18 | Tools case quantity 2 consist of all tools and accessories needed to cutting operation such as trimming knife , knife series , scalpel series , etc. |  |  |

| Item 19 | Microtome, Cryostat | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Floor standing. |  |  |
| 3 | Microtome type: Cross roller bearing rotary and/or rotary retracting. |  |  |
| 4 | Cutting range (µm): to cover the range 2-60 or wider. |  |  |
| 5 | Electronic specimen advance system. |  |  |
| 6 | Digital display for at least: section thickness /actual and preset temperature for chamber. |  |  |
| 7 | Advance range /increments (µm) : 0-10 / 1 or less , 10-20 /2 or less , 20 –60 /5 or less. |  |  |
| 8 | Anti-roll device included. |  |  |
| 9 | Knife guards included. |  |  |
| 10 | Refrigeration: either one or two compressors. |  |  |
| 11 | Compressors must be CFC-free. |  |  |
| 12 | Chamber temperature range Cº: Ambient to (- 35) Cº or wider. |  |  |
| 13 | Defrost: both automatic and programmable. |  |  |
| 14 | Illumination means: to be specified. |  |  |
| 15 | Microtome with universal blade holder to accept high and low profile disposable microtome blades. |  |  |
| 16 | Two sets of specimen holders to be included with the offer. |  |  |
| 17 | Complete start-up kit including consumables and solutions to be included contain at least: Not less than one bottle of lubricating oil. Not less than one bottle of embeddings water soluble matrix for frozen section. |  |  |
| 18 | A list of standard accessories for the offered model |  |  |

| Item 20 | Fully Automated Bench Top Slide Stainer | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Fully automated heavy-duty liner slide Stainer to be programmed for H&E routine stain or other special stain. |  |  |
| 3 | Multi staining capability with load and unload mode drawers system that can perform H&E, PAP and special staining protocols, alone or in combination. |  |  |
| 4 | At least 26 processing stations with at least 20 staining Station. |  |  |
| 5 | At least 450mL Reagent container volume |  |  |
| 6 | At least 30 slide Loading capacity per run |  |  |
| 7 | At least 15 program protocol. Programmable staining time |  |  |
| 8 | Selectable incubation periods |  |  |
| 9 | Dips agitation process. |  |  |
| 10 | Throughput 300Slide /hour |  |  |
| 11 | Water save feature. |  |  |
| 12 | At least one integrated oven/slide dryer for one rack with 30°C - 65°C |  |  |
| 13 | Complete with its necessary accessories and not less than 6 slide racks. |  |  |
| 14 | Fume extraction system, integral fume filter for human and environment safety |  |  |
| 15 | Walk away system. |  |  |

| Item 21 | Microscope, Double Head with Digital Camera and LCD screen | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | 7-fold revolving objective nosepiece, rear facing. |  |  |
| 3 | Semi-Apochromatic infinity objectives; (2x or 2.5x), (4x or 5X), 10x, 20x, 40x , (60x or 63x) and (100x). All objective to be of the same class for use with up to 22 mm field of view |  |  |
| 4 | Face to face Double Head with LED pointer with two colours. |  |  |
| 5 | For the first observer: Trinocular tube with 3 switching positions: 100%/0%, 50%/50%, 0%/100%, viewing angle 30°, F.O.V. up to 22mm, Inter-pupillary distance: 55 to 75mm |  |  |
| 6 | For the second observer: binocular tube, viewing angle 30°, F.O.V. up to 22mm. Inter-pupillary distance: 55 to 75mm |  |  |
| 7 | Two Pairs of adjustable eyepieces 10x/22mm. With IPD adjustment. |  |  |
| 8 | Coaxial double knob focusing and adjustable height stop |  |  |
| 9 | Transmitted light illumination with collector and halogen lamp not less than 30W or LED illumination. |  |  |
| 10 | Condenser: Swing-out with aperture iris diaphragm with variable Koehler illumination. |  |  |
| 11 | Mechanical stage with wide scanning area not less than 75 x 50mm. With slide holder for up to two slides. |  |  |
| 12 | Simple polarizing set. |  |  |
| 13 | Filters: Standard set of filters |  |  |
| 14 | Adjustable light intensity |  |  |
| 15 | All necessary standard accessories and spares to be included. |  |  |
| 16 | A list of additional accessories for the offered model |  |  |
| 17 | A-Digital colour camera : Operation by connection to both; PC via USB (PC Mode) and standalone by connection to a high definition LED Screen (HD Mode) for live image |  |  |
| 18 | B-Capturing JPG images on PC or SD card with max. 2592x1944 pixel, 5Mpixels. |  |  |
| 19 | C-Capturing MP4 movie clips on SD card with max. 1920x1080 pixels, 2Mpixels |  |  |
| 20 | D-Camera control via computer (PC mode) or via remote control (HD mode) |  |  |
| 21 | E-CCD or CMOS sensor (1/2)" |  |  |
| 22 | F-Includes: Operation manual, microscope adapter, USB2 cable, HDMI cable, Power Supply, remote control, SD card, software for use with PC, |  |  |
| 23 | Not less than 48” HD LED Screen, locally supply |  |  |

| Item 22 | Microscope, Fluorescent | Qty. (1) |
| --- | --- | --- |

IMPORTANT NOTE:

Vendors are required to fill out the product details, compliance information, and brochure reference page number in the columns below

TECHNICAL SPECIFICATIONS:

The unit must meet or exceed the requirements listed in the table below.

| Product Details |  |
| --- | --- |
| Name of Manufacturer |  |
| Model/ catalogue number |  |
| Country of Origin for the offered model |  |
| Country where the manufacturer is based |  |
| Delivery time |  |
| Full warranty period |  |

|  | Minimum Requirements | Compliance )Y/N) , Notes | Brochure Page No. |
| --- | --- | --- | --- |
| 1 | The latest highest end model of the required range. |  |  |
| 2 | Fluorescent Illumination with light stop |  |  |
| 3 | 6-fold revolving objective nosepiece, rear facing. |  |  |
| 4 | Four different Exciting Light Filters all installed inside the microscope, at least one for DAPI and one for FTIC |  |  |
| 5 | Ordinary Light System Halogen or LED with adjustable light intensity |  |  |
| 6 | Metal Halide lamp with one spare lamp |  |  |
| 7 | Fluorescent semi-apochromatic infinity Objective: 4X or 5X, 10X, 20X, 40X, (60x or 63x) and 100X oil |  |  |
| 8 | Ergonomic binocular tube, F.O.V. 22 mm, with suitable eyepiece, with tilting angle of 8-32o and telescope mechanism 15-mm movement |  |  |
| 9 | Mechanical stage with wide scanning area of not less than 70x50 mm. |  |  |
| 10 | Pairs of adjustable eyepieces 10x / 22mm. With IPD adjustment |  |  |
| 11 | Nosepiece: Six-tuple coded |  |  |
| 12 | Motorized condenser for bright field (upgradable to dark field, phase contrast and Pol.) with motorized top lens with automatic Koehler elimination. |  |  |
| 13 | A list of standard accessories for the offered model |  |  |

SPECIAL TERMS

- Offers not complying with any of the special terms or the technical specifications shall be considered non-conforming with tender requirements.

- Any vendor providing FORGED documents shall be disqualified from the current tender and banned from participating in any future RMS tenders.

- Wherever term “based” is mentioned it refers to the country where the manufacturing company is founded & established.

- All equipment must be the most recently released model/version which is equal to or higher than the range of the specifications of the required system (low, mid or high) and equal to or higher than the level of technology and required options mentioned in  the technical specifications.

- Required certificates (must be submitted with the technical offer):

- FDA clearance (510K) for equipment of USA based companies.

- MDR (EU) 2017/745 certificate for equipment of EU (European Union) based companies.

- UKCA certificate For Equipment of Great Britain based companies (England, Scotland, Northern Ireland, and Wales).

- ARTG (Australian Register of Therapeutic Goods) certification or approval for Australian and New Zealand based companies.

- Establishment License with the relevant Device License issued by the Therapeutic Products Directorate for Equipment of Canadian based companies.

- PMDA (Pharmaceuticals and Medical Devices Agency) certification or approval for equipment of Japanese based companies.

- Swissmedic (Swiss agency for therapeutic products) certification or approval for equipment of Swiss based companies.

- Norwegian Medicines Agency certification or approval for equipment of Norwegian based companies.

- Only for class Ι medical equipment manufactured by companies based in one of the countries mentioned above, submission of either one of the certificates mentioned above or a free sale certificate in any of these countries shall be accepted.

- For the following equipment:

- All X-ray equipment, MRI, ultrasound, and nuclear medicine systems (regardless where the manufacturing companies are based).

- Where the manufacturing companies are based in other origins than the mentioned in terms 2.1 - 2.8.

The following are required:

- At least two of the certificates mentioned above, one of which has to be FDA clearance (510K) (Only for class Ι medical equipment submission of certificates mentioned in 2.9 shall be accepted).

- Evaluation certificate from the Royal Medical Services for the same offered model with at least 80% passing grade.

If the evaluation is not applicable (based on purchasing committee perspective) bidder should submit a list of installation basis of the same offered model and/ or previous models in at least two of the following hospitals (King Hussein Cancer Center, National Center for Diabetes Endocrinology and Genetic Diseases, Jordan University Hospital or King Abdullah University Hospital) with at least three years of operation, list should include: Name of hospital, Model installed, Quantity, and date of installation.

The purchasing committee has the right to officially contact any of these hospitals and disqualify any offer where the feedback is negative in operation, after sales service or local agent performance.

- The vendor is responsible to ensure through official  documents that classified medical devices are manufactured in conformity with applicable quality system standards (ISO, IEC); (the international quality systems standard for medical device manufacturing ISO 13485) or ISO9001.

- With each offer, bidders must provide a formally endorsed document issued by the manufacturer stating that the bidder is the sole certified agent for the offered item.

- In all of the above cases (except 2.12) certificates must be formally endorsed by JFDA.

- Any vendor not submitting all required certificates will be eliminated.

- Official website of the manufacturer must be provided. The website must clearly demonstrate the history profile and manufacturing features of the company including the same offered model and its brochure.

In case of inconsistency between the information on the website with the provided information of the local agent regarding any offered product, the committee has the right to eliminate that product from the awarding process.

- Offered items should be from reputable well known manufacturers and excellent experience in the field and shall have multiple installations of the same offered model and/ or previous models in RMS main hospitals with at least two years of operation and excellent experience in operation, after sales service & local agent performance; otherwise the purchasing committee has the right to request any of the following:

- An evaluation certificate as mentioned in term 2.10.b.

- A sample of the same offered model at any time during the purchasing process, the sample unit shall be delivered within three weeks from the date of a written notification; any vendor rejects to deliver a sample unit will be eliminated. Evaluation duration will be determined by the purchasing committee.

Any offered item fail in the evaluation/assessment process will be rejected

- Vendors must specify the origin of the offered items and accessories in the technical offer.

- Bidders must submit their reservations/queries regarding tender specifications and/or special terms within the first half of the tender closing period starting from the tender announcement date. Reservations/queries submitted after the end of this period shall be rejected.

- Goods are to be brand new; manufacturing date of the delivered products should not exceed 18 months from the date of the final award.

- Delivery period should be mentioned clearly; the purchasing committee has the right to reject any offer with delivery time exceeding 4 months.

- Warranty:

- Offers must include a full warranty for a period of a minimum of 24 months from the date of installation or 30 months from the date of receiving the items at the agreed location mentioned in the final order (whichever comes first).

Warranty must include corrective & preventive maintenance activities as per manufacturer recommendations including:

- Required spare parts (free of charge)

- Labour

- Hardware

- Software

- Rechargeable batteries

At the end of the warranty period, The supplier commits to implement final inspection of the submitted goods and submit the reports signed by the site chief engineer stating that the equipment are working properly as well as all preventive maintenance reports during the warranty period.

Warranty & installation are excluded for items mentioned in Attachment no.1 the purchasing committee has the right to exclude any other item not listed in the attachment and does not required installation and warranty.

- In the case where a delay in installation has occurred as a result of the supplier's dereliction and has exceeded a period of one month from the date of receiving the items, the approved warranty period will automatically be considered as 24 months from the date of installation.

- If at any time during the warranty period the item becomes inoperative due to a technical fault the item must then be repaired by the supplier /local agent within a period of 5 working days from written notification, warranty will be extended according to downtime period.

- If the delay exceeds 30 days the supplier must replace the item with a new identical functioning one (within the same delivery period mentioned in the final offer). In case the item was replaced by a new one, the warranty period mentioned in 9.a) above will start from the installation and commissioning date of the new item.

- Local agent/ supplier is committed to transport and install any awarded item during the warranty period on free of charge basis to any location inside the country whenever required by Royal Medical Services, this should not include pre-installation requirements in the new site.

- All offered items (main unit) should be fully designed, manufactured, and labelled by their real original manufacturer in which all related testing, research, development and approvals went through.

Any relabelled products for the main unit (white-label manufacturing, OEM, or repackaging) are rejected.

- One set of operation manual(s) and one set of service manual(s) including schematics and a spare-part list must be delivered with each unit, CD/DVD is acceptable. For large tenders, a certain agreed percentage of manuals per item may be agreed upon.

- Where applicable, pre-installation shall be the sole responsibility of the supplier. Pre-installation shall include removal of old system(s), any civil work, electrical work or site modification(s) necessary to accommodate the new system(s) according to manufacturers’ specifications and safety standards in addition to the work required for bringing back the site to the same working conditions as before installing the new system(s).

- Power requirements: where applicable either single phase 220V, 50Hz or 3-phase 380V. Systems with external transformers are considered conforming only if clearly stated in the technical specifications.

- Technical offers must include clear original technical brochures/catalogues for all offered items.

- Offers must include fully detailed technical offers and compliance sheets as a soft copy ( either Microsoft office or Microsoft excel format) in addition to a hard copy, mentioning the exact model/catalogue number and country of origin of the offered item(s), full technical description/specifications and any accessories or options included in the offer.

Offer must clearly indicate the origin of the offered item and the country where the manufacturer of the offered items is based.

- Compliance sheets must be as per the tabular format of the technical specifications in the tender documents, listing the required specifications on one column and a  Yes  or NO  response to each point in the adjacent column, with reference to page and line numbers in the relevant technical brochure. Offers not complying with this term shall be rejected.

- Qualifications and after sales service: The technical bid should contain all the necessary documents (training certificates) to prove the capability of the bidders for supplying and installing a trouble free equipment meeting the quality standards and technical specifications of the manufacturer and the ability of the bidders for providing efficient after sales service conducted by authorized certified biomedical service engineers with minimum 2 years experience in the same field.

- Accessories and consumables:

  - Any accessories and consumable items necessary to operate the offered system must be clearly identified and priced separately.

All offered accessories and consumables must be approved by the manufacturer.

  - Technical offers must include a priced list for all accessories and consumables related to the required equipment as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) with prices fixed for a period of five years from the date of installation and commissioning with a maximum annual increase of 2%, any essential item not listed will be considered free of charge.

  - Accessories and consumables must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Where applicable, a start-up kit of accessories and consumable items must be provided with each system on a free-of-charge basis.

- Spare Parts:

  - Technical offers must include a comprehensive and priced list of all spare parts related to the awarded equipment (including rechargeable batteries) as a hard copy in addition to a soft copy (either Microsoft office or Microsoft excel format) valid for a minimum period of five years with a maximum annual increase of 2%, commencing at the end date of the warranty period, any essential item not listed will be considered free of charge.

  - Spare parts must be priced according to their delivery destination either to Queen Alia International Airport or to RMS Main Medical Stores.

  - Prices of spare parts should be reasonable and will be taken into consideration during the purchasing process; the purchasing committee has the right to eliminate any offered item with unreasonable high prices of spare parts.

  - Delivery period of required spare parts should not exceed 2 months from the date of the final order.

- Spare parts, consumables and accessories availability must be guaranteed for a minimum period of ten years starting from the date of installation and commissioning.

- Local agent is committed to sign a service contract with Royal Medical Services to perform all relevant maintenance for the awarded equipment (whenever required by DRMS within a period that does not exceed one year from the end date of the warranty period and if mentioned clearly in the specifications) without exceeding the original mentioned cost and as per mentioned terms within minimum required period.

- Tender Awards:

  - For the final list of offers having a chance of winning the award, the awarding process shall be based on the accumulative value of both the offered item and its’ running cost (Total Cost of Ownership) over a period of seven years from the date of installation and commissioning. Only offers with the lowest total cost of ownership over a period of seven years from the date of installation and commissioning shall qualify for the award.

  - Running cost includes the value of consumables, accessories needed to operate the system over the same period as well as the cost of any service contract (where applicable).

- For PC/Laptop based systems:

  - Complete restoration medium (CD/DVD/etc.) of the operating system and the application software must be supplied.

  - Purchased licenses, software keys & dongles must be provided at any time during and after the warranty period on free of charge basis whenever required by DRMS for a minimum period of ten years starting from the date of installation and commissioning.

  - Where locally supplied computers, laptops & printers are offered, the offered model should be from well known manufacturer.

- Pricing must include services of sale, shipment, transportation, delivery from port to site or to Main Medical Stores, installation, pre-installation (if needed), training, commissioning, warranty and bringing the equipment into service.

- Custom clearance of goods shall be the responsibility of the Jordanian Armed Forces (JAF), however, suppliers shall bear all costs incurred by handling charges and any demurrage charges or extra expenses incurred by the port’s corporation (including expenses caused by delay in presenting the necessary shipment documents for either clearing or transporting the goods to the required location mentioned in the final order, delivery note issuing charges, unloading charges, local shipping charges etc.). The supplier is also responsible for providing of all relevant shipping documents, together with the delivery order(s).

- a. DRMS has the right to increase the awarded quantities by a percentage not exceeding 35% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

b. DRMS has the right to decrease the awarded quantities by a percentage not exceeding 50% after the final order notification with the same prices, terms and conditions of the original contract upon DRMS request and approval of the awarded party.

- a. Bidder is not allowed to submit more than one offer for the same item whether that’s solely or in coalition or partnership with other bidders.

b. Bidder is allowed to include within their offer optional alternatives for the same offered item from the same manufacturer.

- The supplier must furnish DRMS with a guarantee stamped and legalized by the Notary Public equals to (115%) of the total value of the awarded equipment valid for twelve months from the date of final acceptance of the equipment by DRMS.

- Training:

  - For items where abroad service training courses are required in technical specifications, offers must include a certified service training program for at least 3 working days at a reputable centre abroad recognized by the manufacturer for at least one biomedical engineer or biomedical technician; all costs inclusive, air tickets,, boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - For items where abroad user  training courses are required in technical specifications, offers must include a certified operator training program at a reputable center abroad recognized by the manufacturer for at least one operator; all costs inclusive, air tickets, boarding, commuting, accommodation (minimum 3 star hotel on full board basis) and any extra costs.

  - The period of the training courses must be according to the manufacturer’s program excluding traveling days and must be stated clearly in the technical offer.

  - Training Programs must conform to the following standards:

- User training must comprise understanding and use of operation manual(s), correct and safe operation of the equipment, as well as user preventive maintenance and calibration.

- Service training must comprise: theory, understanding and use of service manual(s), calibration, preventive maintenance procedure, and practical troubleshooting and repair exercises, and must be conducted by professional instructors employed or authorized by the system manufacturer.

- Service training must be conducted on a system of identical make, model, and configuration to that purchased by DRMS, and designated by the manufacturer or the local agent for training purposes.

- Certificates must be endorsed and officially sealed by the system manufacturer, legally empowering trainees to engage in user and service activities according to operation and service manual(s).

- Where applicable, offers must include an on-site user and service training.

30. For offers submitted in Jordanian dinars, Payment will be either by wire transfer or by cheque after final acceptance of goods. Any other way of payment will be rejected.

Attachment 1

| # | Equipment |
| --- | --- |
| 1 | Adjustable Auto Pipettes |
| 2 | Resuscitation Bag |
| 3 | Laryngoscope Set |
| 4 | Oxygen Flow meter wall type/ single |
| 5 | Regulator Suction with canister, wall vacuum outlet |
| 6 | Oxygen Regulator for Oxygen Cylinders |
| 7 | Pulse Oximeter, Finger type |
| 8 | Oxygen Cylinder |
| 9 | Doppler, portable |
| 10 | Diagnostic set, Portable |
| 11 | Direct Ophthalmoscope, Portable |
| 12 | Otoscope, Portable |
| 13 | Air Mattress System, homecare |
| 14 | Stethoscope |
| 15 | Aneroid Sphygmomanometer |
| 16 | Video Assisted Laryngoscope, portable |
| 17 | Scale, Manual |
| 18 | Wood's Light |
| 19 | Cough Pressure, Normal Saline |
| 20 | Rehabilitation Walking Parallel Bars, non-powered |
| 21 | Therapy Mat |
| 22 | Medical Ball All Size |
| 23 | Dumbbells Rack with complete set of dumbbells |
| 24 | Crutches |
| 25 | Shoulder wheel |
| 26 | Mobile Mirror |
| 27 | Cuff Weights |
| 28 | Walker, different sizes |
| 29 | Patient Elbow Stick |
| 30 | Rehabilitation Training Ladder |
| 31 | Rehabilitation Suspension Frame |
| 32 | Exercise Band All Size (Theraband) |
| 33 | Lens trial set |
| 34 | Wheel Chair |
| 35 | Commode Chair |
| 36 | Bassinet (Baby Cot) |
| 37 | Resuscitation Cart (Crash Cart) |
| 38 | Medication Cart |
| 39 | Cart, Drawers |
| 40 | Examination Couch, Manual |
| 41 | Gynaecology Examination Table, Manual |
| 42 | Examination Table, Neonates, Manual |
| 43 | Intravenous Pole, Mobile Stand |
| 44 | DDA Cabinet |
| 45 | Stainless steel Multipurpose Trolley |
| 46 | Cabinet, Instrument, Operation Theatres |
| 47 | Dressing Cart |
| 48 | Stainless steel wire shelving unit |
| 49 | Paper Trolley |
| 50 | Stainless Steel Sink (Clean up counter) |
| 51 | Scopes Cabinet |
| 52 | Mayo Table |
| 53 | Table, Instrument |
| 54 | Stool, Adjustable, Doctor |
| 55 | Stool, Adjustable, Operation Theatres |
| 56 | Carts, linen/laundry, soiled, Double |
| 57 | Closed distribution trolley |
| 58 | Step Ladder, Conductive, Double |
| 59 | Step, Surgeon, Single |
| 60 | Kick Bucket |
| 61 | Mobile Stand for Oxygen Cylinder |
| 62 | Stainless Steel Wire Basket, 1 STU |
| 63 | Cart, Plaster |