| Specification | Parameter | No. |
| --- | --- | --- |
| OPTICS |  |
| Simultaneous |  |
| 160-900 | Full Wavelength Coverage/nm | 1 |
|  | Automatic wave length selection | 2 |
| <0.5 | Wavelenght accuracy / nm | 3 |
| 0.002-0.007 | Optical resolution ( pm/nm) | 4 |
|  | Purged optics | 5 |
| Echelle | Monochromator | 6 |
| (no purge gas required) | Vaccum spectrometer | 7 |
| Detector |  |
| Peltier-cooled (-10 ºC) CCD |  | 8 |
| Fast simultaneous background correction |  | 9 |
| Neon | Wavelenght calibration | 10 |
| 1840*1840 pixels | Detector resolutin | 11 |
| PLASMA |  |
|  | Vertical-orientated plasma | 12 |
|  | Horizontal-orientated plasma | 13 |
|  | Radially-viewed configuration | 14 |
|  | Axially-viewed configuration | 15 |
|  | Dual view configuration | 16 |
|  | Shear gas (plasma tail removal) | 17 |
|  | Extended outer tube | 18 |
|  | Demountable torch | 19 |
| GENERATOR |  |
| Free running | Type of generator | 20 |
| 27 MHz | Generator frequency | 21 |
|  | Variable power | 22 |
| 700 | Minimum power / W | 23 |
| 1700 | Maximum power / W | 24 |
| 0-20 | Plasma gas Argon flow/L min | 25 |
| 0-20 L/min during analysis | Argon gas flow rate | 26 |
| 99.95% | Argon gas purity for ignition | 27 |
| <20 | Average total consumption/L min | 28 |
|  | Eco mode to reduce the plasma gas flow rate | 29 |
| SAMPLE INTRODUCTION |  |
|  | Concentric nebulizer | 30 |
|  | Cyclonic spray chamber | 31 |
|  | Fast sample change | 32 |
| SOFTWARE |  |
|  | Software optimization of plasma | 33 |
|  | Software optimization of observation position | 34 |
|  | Multivariate evaluation ( incl.IEC) | 35 |
|  | Automated multi calibration (using diff. Lines) | 36 |
|  | Line library ( for multivariate evaluation) | 37 |
|  | Automatic system shutdown after analysis | 38 |
| AUTO SAMPLER |  |
|  | Auto diluent auto sampler with different position for standards and samples -50 ml tubes, not less than 10 positions. -15ml tubes, not less than(60/1-rack),3 pieces | 39 |
|  | Automated oxygen ashing | 40 |
|  | Full desktop computer | 41 |
|  | Color laser jet printer | 42 |
| ENVIRONMENTAL REQYUIREMENTS |  |
|  | External water cooling system | 43 |
| 230 V ±10% ; 50Hz, | Power requirements | 44 |
| External water cooling system | Relative Humidity | 45 |
| +10 ºC up to 35 ºC | Temperature | 46 |
| USA or Europe | Origin | 47 |
|  | Low gas consumption | 48 |
|  | Duct venting | 49 |
| ADDITIONAL |  |
| Capability of software to distinguish between the overlapping elements | 50 |
| Possibility of adding an element and wavelength without re-measurements. | 51 |
| Complies with iso 17025:2017 | 52 |
| Microwave digestion system with enclosed specifications | 53 |

Microwave digestion System

Features:

•   Round, pressure resistant oven chamber made of PFA-coated stainless steel with increased safety and security.

•   Router with at least 6  pressure  vessels  that can operate at 400 psi and 250°C simultaneously

•   Temperature range up to 250°C

•   Sample weight ranges from 0.1 to 10 g

•   Microwave power: at least 1200.W

•   Minimum vessel volume 60 - 100 ml

•   Temperature sensors with thermowell for the control vessel digestion vessels for temperature control.

•   Additional temperature measuring and control for all vessels

•   Direct sensing and control of sample temperature

•   Flexibility to program for specific customizable sample preparation and automated recording of conditions during each digestion.

•   Digital display for viewing and programming the parameters, OR the bidder has to offer a dedicated software and PC to run and control the system.

•   safety interlocks

•   Fume exhaust module

•   Suitable venting system for vessel when exceeding pressure limits coupled with  detection mechanisms for automatic shuts down

•   Teflon lining of oven cavity

•   Digestion vessels molded with high chemical and resistant material

•   Release of acid fumes into collection reservoir in vessel carousel

•   Automatic pressure release mechanism in oven cavity

•   Power supply: 220V 50Hz, MK plug.

Microwave digestion System

•   All required accessories for optimum operation shall be included in the offer.

•   Operation and service manual has to be delivered along with equipment.

•   Onsite training should be provided.

•        All components should be priced separately to be able to choose the best combination of options and accessories in terms of the number of vessels and number and size of rotors etc.

**Specifications for kjeldahl nitrogen analyzer

*. Digestion unit:

-- fit for 20samples

-- exact temp/time control

-- scrubber unit for gases and fumes

-- tubes for digestion(100tube)(300ml)

-- over heating protection

-- providing antifoam and catalyst

*. Distillation unit:

-- programmed and automatic addition of H3PO4 ,H2O, NaOH.

-- programmed and automatic suction of wast and receiver.

-- connection to computer with software.

-- print out for result.

-- tanks for H3PO4 ,H2O, NaOH,wast.

-- automatic titration.

-- computer and printer.

-- pH electrode system(kcl or gel electrolyte)

* optional:-Auto sampler for distillation.

* provide certified reference material.

/2017    * complies with ISO 17025

الشروط الخاصة

1. يتم تركيب الجهاز والتدريب عليه من قبل الشركة الموردة في المختبرات العسكرية لمراقبة الجودة.

2. تدريب داخلي للفنيين في المختبر ولمدة لا تقل عن أسبوع.

3. يتم التدريب الخارجي لأجهزة ( ICP-OES, Kjeldahl nitrogen analyzer) لثلاثة أشخاص من الفنيين لكل جهاز ولمدة (4) أربعة أيام في بلد المنشأ على نفقة الشركة المحال عليها بعد (3) ثلاثة أشهر من إنتهاء التدريب الداخلي.

4. كفالة الجهاز لمدة (3) أعوام من تاريخ التشغيل.

5. جميع الأسعار خاضعة لنسبة ( الصفر) من قانون الضريبة العامة على المبيعات.

6. التسليم واصل مستودعات المختبرات العسكرية لمراقبة الجودة.

7. تلتزم الشركة بتقديم شهادة معايرة لجميع الأجهزة.

8. جميع الأجهزة منشأ ( أوروبا الغربية, الولايات المتحدة الأمريكية).

9. يجب وجود وكيل معتمد للشركة الصانعة داخل المملكة على أن يتوفر لديهم مهندس معتمد حاصل على شهادة معتمدة من الشركة الصانعة 0

10. يلتزم المتعهد بصيانة الجهاز صيانة مجانية لمدة (3) ثلاث سنوات على الاقل تبدأ من تاريخ الاستلام النهائي للجهاز شاملة لقطع الغيار, وتكون قطع الغيار المستبدلة اصلية وجديدة 100% .

11. الالتزام بتقديم قائمة بأسعار قطع الغيار و المستهلكات اللازمة للجهاز والتعهد بها لمدة لا تقل عن خمس سنوات بحيث تكون هذه الأسعار ثابتة بعد انتهاء مدة الكفالة وهي ( 3 ) سنوات والتعهد بتوفير قطع غيار للجهاز لمدة لا تقل عن عشر سنوات.

12. يلتزم المتعهد بأن تكون اللوازم الموردة جديدة ( 100%) وخالية من أية عيوب في الصنع أو في المادة أو من طراز حديث وأن تكون سنة الصنع ( 2019) فما فوق.

13. يلتزم المتعهد بتوفير تدريب متخصص على برامج وطريقة عمل الجهاز والصيانة الوقائية وعملية تبديل القطع لكل جهاز.

14. يلتزم المتعهد أن يرفق مع الجهاز دليل الصيانة وتعليمات التشغيل وبرمجيات الجهاز.

15. في حال توقف الجهاز عن العمل تلتزم الشركة بتقديم الصيانة الفورية خلال (13) يوم من تاريخ التبليغ وفي حال تجاوز هذه المدة يفرض عليه غرامة ماليه مقدارها 5% من قيمة الصيانة 0

16. يلتزم المتعهد باستمرارية تحديث وتطوير ال software التابع للجهاز مجاناً وبشكل دوري 0

17. أن يكون الجهاز موديل 2019 أو 2020 0

# Text boxes

Inductively Coupled Plasma (ICP)